## 0425S
Below is Port 1 RF Low Level Return Loss - mirror of RF Return Loss | Below is Port 4 with insufficient Power
Above Data with A0030 Buffer Amplifier and Normalized Reference
Above PNA Data with no Amplifier - Not Used

## Mapping
Avoid using '-' in sheet names
Sheet names must be spelled correctly
Axis column labels can be upper or lowercased.
Number of lines must be indicated
Conversion Loss | [8 empty] | Input IP3 | [8 empty] | LO to RF Isolation
Lines | 2 | [7 empty] | Lines | 2 | [7 empty] | Lines | 2
Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max | Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max | Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max
Config A | CL RFRL | A | 5 | 205 | CL RFRL | C | 5 | 205 | Config A | IP3 | A | 5 | 103 | IP3 | C | 5 | 103 | Config A | LORL LRI RI | V | 5 | 205 | LORL LRI RI | X | 5 | 205
Config B | CL RFRL | A | 5 | 205 | CL RFRL | F | 5 | 205 | Config B | IP3 | A | 5 | 103 | IP3 | F | 5 | 103 | Config B | LORL LRI RI | V | 5 | 205 | LORL LRI RI | AB | 5 | 205
LO to IF Isolation | [8 empty] | RF to IF Isolation | [8 empty] | IF Response
Lines | 2 | [7 empty] | Lines | 2 | [7 empty] | Lines | 2
Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max | Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max | Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max
Config A | LORL LRI RI | V | 5 | 205 | LORL LRI RI | Z | 5 | 205 | Config A | LORL LRI RI | V | 5 | 205 | LORL LRI RI | AD | 5 | 205 | Config A | IFCL IFRL | A | 3 | 103 | IFCL IFRL | D | 3 | 103
Config B | LORL LRI RI | V | 5 | 205 | LORL LRI RI | AD | 5 | 205 | Config B | LORL LRI RI | V | 5 | 205 | LORL LRI RI | Z | 5 | 205 | Config B | IFCL IFRL | A | 3 | 103 | IFCL IFRL | H | 3 | 103
Conversion Loss vs. LO Power | [8 empty] | Input IP3 vs. LO Power | [11 empty] | A Configuration Downconversion
Lines | 4 | [7 empty] | Lines | 0 | [7 empty] | -10 dBm RF Input | 0xLO | 1xLO | 2xLO | 3xLO | 4xLO | 5xLO
Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max | Label | Sheet | X axis | Cell Min | Cell Max | Sheet | Y axis | Cell Min | Cell Max | 1xRF
+22 dBm | CL vs. LO | A | 4 | 204 | CL vs. LO | C | 4 | 204 | [9 empty] | 2xRF
+17 dBm | CL vs. LO | A | 4 | 204 | CL vs. LO | D | 4 | 204 | [9 empty] | 3xRF
+15.5 dBm | CL vs. LO | A | 4 | 204 | CL vs. LO | E | 4 | 204 | [9 empty] | 4xRF
+14.5 dBm | CL vs. LO | A | 4 | 204 | CL vs. LO | F | 4 | 204 | [9 empty] | 5xRF

## CL RFRL
RF (GHz) | Standard "A" CL | Standard "A" RF R.L. | Standard "B" CL | Standard "B" RF R.L. | 1-dB Compression
| -8.1978 |  | -9.0211
1 | -74.2528 | -3.6951 | -26.6334 | -3.4412 | Std "B" | 14.5
1.1275 | -72.2925 | -3.5281 | -25.4644 | -3.3838
1.255 | -71.0369 | -3.3354 | -23.7597 | -3.3113
1.3825 | -65.3183 | -3.125 | -21.9152 | -3.2243
1.51 | -62.7135 | -2.9056 | -20.5876 | -3.131
1.6375 | -58.2954 | -2.6777 | -19.5324 | -3.0273 | Std "A" | 14.5
1.765 | -51.4238 | -2.4989 | -18.5989 | -2.9246
1.8925 | -43.2799 | -2.3695 | -18.1046 | -2.8261
2.02 | -38.3673 | -2.294 | -17.7504 | -2.733
2.1475 | -29.6161 | -2.2453 | -17.3487 | -2.638
2.275 | -23.9624 | -2.2099 | -16.9703 | -2.5498
2.4025 | -20.7174 | -2.1997 | -16.763 | -2.4732
2.53 | -18.0869 | -2.2108 | -16.4536 | -2.4118
2.6575 | -16.5267 | -2.2164 | -16.0223 | -2.3546
2.785 | -15.6201 | -2.2456 | -15.4677 | -2.3211
2.9125 | -15.6038 | -2.2889 | -14.8721 | -2.3022
3.04 | -15.3785 | -2.3423 | -14.0418 | -2.3226
3.1675 | -14.8324 | -2.3818 | -13.3991 | -2.3656
3.295 | -14.2475 | -2.4807 | -12.5781 | -2.4324
3.4225 | -13.852 | -2.6251 | -11.8934 | -2.5292
3.55 | -12.6883 | -2.7839 | -11.3224 | -2.6984
3.6775 | -11.6769 | -2.962 | -11.0452 | -2.9451
3.805 | -10.9914 | -3.202 | -10.5157 | -3.278
3.9325 | -10.5321 | -3.458 | -10.1377 | -3.7311
4.06 | -9.8298 | -3.7285 | -9.8269 | -4.3264
4.1875 | -9.6088 | -4.0229 | -9.3107 | -5.0687
4.315 | -9.5254 | -4.3737 | -8.6642 | -5.9865
4.4425 | -9.3448 | -4.7211 | -8.2534 | -7.1239
4.57 | -9.0296 | -5.0825 | -7.985 | -8.4269
4.6975 | -8.8288 | -5.5028 | -7.6988 | -9.7055
4.825 | -8.5988 | -5.9855 | -7.5761 | -10.8653
4.9525 | -8.3284 | -6.5038 | -7.5389 | -11.8867
5.08 | -8.0553 | -7.0643 | -7.5318 | -12.6098
5.2075 | -7.8616 | -7.7013 | -7.5805 | -13.0391
5.335 | -7.7125 | -8.4086 | -7.6303 | -13.1719
5.4625 | -7.5464 | -9.1521 | -7.6834 | -13.0281
5.59 | -7.4087 | -9.9376 | -7.7722 | -12.5809
5.7175 | -7.3102 | -10.7701 | -7.8905 | -11.8884
5.845 | -7.2218 | -11.5881 | -7.95 | -11.1278
5.9725 | -7.1346 | -12.501 | -7.9897 | -10.3745
6.1 | -7.0724 | -13.4842 | -8.0107 | -9.6686
6.2275 | -7.0141 | -14.6032 | -8.0077 | -9.1339
6.355 | -6.9522 | -15.7997 | -7.9687 | -8.7513
6.4825 | -6.9025 | -17.0331 | -7.928 | -8.5001
6.61 | -6.8549 | -18.2843 | -7.8928 | -8.3483
6.7375 | -6.8149 | -19.6261 | -7.8668 | -8.2795
6.865 | -6.8012 | -20.9625 | -7.843 | -8.2717
6.9925 | -6.8057 | -22.2619 | -7.8213 | -8.2913
7.12 | -6.8245 | -23.2696 | -7.815 | -8.343
7.2475 | -6.8514 | -24.0093 | -7.8165 | -8.383
7.375 | -6.8905 | -24.4156 | -7.8361 | -8.3819
7.5025 | -6.9481 | -24.5906 | -7.8767 | -8.3607
7.63 | -7.0262 | -24.4864 | -7.9557 | -8.2895
7.7575 | -7.1007 | -24.0602 | -8.049 | -8.1646
7.885 | -7.186 | -23.2672 | -8.1597 | -7.9882
8.0125 | -7.275 | -22.2534 | -8.2912 | -7.7689
8.14 | -7.3768 | -21.0669 | -8.4568 | -7.5167
8.2675 | -7.4603 | -19.8756 | -8.642 | -7.2487
8.395 | -7.5645 | -18.7011 | -8.8341 | -6.9844
8.5225 | -7.6563 | -17.5522 | -9.0239 | -6.7439
8.65 | -7.7408 | -16.3751 | -9.1824 | -6.5217
8.7775 | -7.8085 | -15.2601 | -9.3099 | -6.3442
8.905 | -7.8686 | -14.2899 | -9.3857 | -6.2152
9.0325 | -7.9131 | -13.447 | -9.4193 | -6.1338
9.16 | -7.9516 | -12.7019 | -9.4198 | -6.1043
9.2875 | -7.968 | -12.0727 | -9.4183 | -6.1283
9.415 | -7.9608 | -11.505 | -9.3768 | -6.1842
9.5425 | -7.9501 | -10.9898 | -9.3072 | -6.2738
9.67 | -7.9378 | -10.5462 | -9.2606 | -6.4038
9.7975 | -7.9324 | -10.125 | -9.2251 | -6.5689
9.925 | -7.939 | -9.7409 | -9.1796 | -6.7814
10.0525 | -7.9684 | -9.3852 | -9.1353 | -7.0395
10.18 | -7.9913 | -9.0717 | -9.0815 | -7.3353
10.3075 | -8.0221 | -8.7642 | -9.0086 | -7.6673
10.435 | -8.0337 | -8.4827 | -8.9275 | -8.0569
10.5625 | -8.0522 | -8.247 | -8.8417 | -8.5099
10.69 | -8.0328 | -8.0457 | -8.7489 | -9.0311
10.8175 | -8.0076 | -7.8654 | -8.676 | -9.5992
10.945 | -7.9635 | -7.7274 | -8.5915 | -10.2254
11.0725 | -7.9212 | -7.6047 | -8.5166 | -10.8925
11.2 | -7.8664 | -7.4783 | -8.4458 | -11.6112
11.3275 | -7.8338 | -7.3461 | -8.3918 | -12.4118
11.455 | -7.8437 | -7.2247 | -8.3603 | -13.2544
11.5825 | -7.8647 | -7.089 | -8.3602 | -14.2441
11.71 | -7.9045 | -6.9322 | -8.3772 | -15.3232
11.8375 | -7.9779 | -6.7576 | -8.4044 | -16.5959
11.965 | -8.082 | -6.5669 | -8.457 | -18.0547
12.0925 | -8.1892 | -6.3507 | -8.5252 | -19.3663
12.22 | -8.2981 | -6.1388 | -8.6124 | -20.3738
12.3475 | -8.4262 | -5.9636 | -8.7102 | -20.8868
12.475 | -8.5062 | -5.8095 | -8.8114 | -21.0366
12.6025 | -8.5443 | -5.6785 | -8.9023 | -20.8882
12.73 | -8.5343 | -5.5766 | -8.9745 | -20.3534
12.8575 | -8.4994 | -5.5058 | -9.0182 | -19.5259
12.985 | -8.4178 | -5.4771 | -9.05 | -18.3308
13.1125 | -8.3632 | -5.4825 | -9.0619 | -16.7819
13.24 | -8.3027 | -5.5246 | -9.0748 | -15.1947
13.3675 | -8.255 | -5.5776 | -9.0743 | -13.755
13.495 | -8.2273 | -5.6313 | -9.1009 | -12.6478
13.6225 | -8.2144 | -5.6996 | -9.1088 | -11.7312
13.75 | -8.2136 | -5.7587 | -9.1506 | -10.9679
13.8775 | -8.2105 | -5.7822 | -9.1841 | -10.3182
14.005 | -8.2246 | -5.8302 | -9.238 | -9.7677
14.1325 | -8.28 | -5.8484 | -9.3053 | -9.2627
14.26 | -8.3153 | -5.8446 | -9.361 | -8.8191
14.3875 | -8.3514 | -5.8425 | -9.4331 | -8.4261
14.515 | -8.4408 | -5.8547 | -9.5203 | -8.0749
14.6425 | -8.5025 | -5.8777 | -9.6143 | -7.7495
14.77 | -8.5086 | -5.9102 | -9.6795 | -7.4453
14.8975 | -8.5432 | -5.9672 | -9.7864 | -7.1846
15.025 | -8.5295 | -6.0853 | -9.8551 | -6.9476
15.1525 | -8.4769 | -6.2043 | -9.909 | -6.7251
15.28 | -8.4164 | -6.3724 | -9.9413 | -6.5464
15.4075 | -8.3666 | -6.576 | -9.989 | -6.3915
15.535 | -8.3048 | -6.8353 | -9.9967 | -6.2816
15.6625 | -8.2798 | -7.1093 | -10.0244 | -6.1953
15.79 | -8.2049 | -7.3778 | -10.0388 | -6.1369
15.9175 | -8.1629 | -7.7062 | -10.0386 | -6.12
16.045 | -8.1495 | -8.0796 | -10.0423 | -6.1222
16.1725 | -8.0677 | -8.4795 | -10.0095 | -6.1541
16.3 | -7.9865 | -8.9258 | -9.9455 | -6.2215
16.4275 | -7.9556 | -9.4378 | -9.8762 | -6.3053
16.555 | -7.8874 | -9.9867 | -9.7963 | -6.4158
16.6825 | -7.7801 | -10.6022 | -9.6633 | -6.5571
16.81 | -7.7629 | -11.2841 | -9.5884 | -6.7184
16.9375 | -7.7068 | -12.0714 | -9.5011 | -6.9151
17.065 | -7.6523 | -12.9296 | -9.4206 | -7.1228
17.1925 | -7.6272 | -13.8632 | -9.3477 | -7.3444
17.32 | -7.588 | -14.9827 | -9.2843 | -7.5932
17.4475 | -7.5564 | -16.3833 | -9.2018 | -7.8557
17.575 | -7.5483 | -17.832 | -9.1338 | -8.124
17.7025 | -7.5239 | -19.9128 | -9.0596 | -8.411
17.83 | -7.5286 | -21.8254 | -9.0077 | -8.7047
17.9575 | -7.5396 | -23.0947 | -8.9485 | -9.0291
18.085 | -7.5237 | -23.8445 | -8.8924 | -9.3609
18.2125 | -7.5401 | -24.2552 | -8.8468 | -9.705
18.34 | -7.5685 | -24.2824 | -8.8023 | -10.0759
18.4675 | -7.5813 | -23.8809 | -8.7631 | -10.4632
18.595 | -7.6309 | -22.9914 | -8.7548 | -10.8555
18.7225 | -7.6825 | -21.8377 | -8.7671 | -11.2525
18.85 | -7.7191 | -19.8929 | -8.7703 | -11.6623
18.9775 | -7.7853 | -17.9047 | -8.8073 | -12.0587
19.105 | -7.848 | -16.3718 | -8.8347 | -12.4442
19.2325 | -7.9037 | -15.1728 | -8.865 | -12.7991
19.36 | -7.9977 | -14.0936 | -8.8939 | -13.1532
19.4875 | -8.1017 | -13.1708 | -8.9549 | -13.493
19.615 | -8.1949 | -12.3712 | -8.988 | -13.7876
19.7425 | -8.3213 | -11.64 | -9.0442 | -14.0572
19.87 | -8.5076 | -10.9812 | -9.1051 | -14.2785
19.9975 | -8.6263 | -10.3739 | -9.1436 | -14.4245
20.125 | -8.76 | -9.8146 | -9.1882 | -14.4826
20.2525 | -8.9 | -9.3078 | -9.2536 | -14.4366
20.38 | -8.9864 | -8.8591 | -9.3096 | -14.35
20.5075 | -9.0278 | -8.4483 | -9.3385 | -14.2034
20.635 | -9.102 | -8.1129 | -9.4048 | -13.9928
20.7625 | -9.1496 | -7.8423 | -9.4325 | -13.7525
20.89 | -9.2299 | -7.5786 | -9.4856 | -13.4261
21.0175 | -9.3424 | -7.322 | -9.5096 | -13.0569
21.145 | -9.4063 | -7.0972 | -9.5413 | -12.6843
21.2725 | -9.4719 | -6.8954 | -9.5436 | -12.3605
21.4 | -9.5413 | -6.7153 | -9.5701 | -12.0895
21.5275 | -9.5556 | -6.5403 | -9.5416 | -11.847
21.655 | -9.5776 | -6.3895 | -9.5439 | -11.6366
21.7825 | -9.6304 | -6.245 | -9.5524 | -11.4611
21.91 | -9.6694 | -6.0949 | -9.5659 | -11.3129
22.0375 | -9.7 | -6.006 | -9.5762 | -11.2825
22.165 | -9.7296 | -5.9546 | -9.6158 | -11.3517
22.2925 | -9.728 | -5.9233 | -9.6334 | -11.4688
22.42 | -9.7175 | -5.9089 | -9.6644 | -11.6601
22.5475 | -9.7151 | -5.9234 | -9.706 | -11.9192
22.675 | -9.7011 | -5.9506 | -9.7405 | -12.2145
22.8025 | -9.6941 | -5.9868 | -9.784 | -12.5945
22.93 | -9.6799 | -6.0405 | -9.8378 | -13.0467
23.0575 | -9.6873 | -6.1108 | -9.9015 | -13.6409
23.185 | -9.6687 | -6.1943 | -9.9539 | -14.2781
23.3125 | -9.6626 | -6.2922 | -10.0288 | -15.0593
23.44 | -9.6397 | -6.4016 | -10.0768 | -16.0638
23.5675 | -9.6208 | -6.5206 | -10.143 | -17.4503
23.695 | -9.575 | -6.6459 | -10.2148 | -19.5387
23.8225 | -9.5722 | -6.7809 | -10.2959 | -22.075
23.95 | -9.5609 | -6.9254 | -10.3878 | -23.3836
24.0775 | -9.5733 | -7.0635 | -10.5288 | -24.1991
24.205 | -9.5888 | -7.2206 | -10.638 | -24.5052
24.3325 | -9.621 | -7.3757 | -10.7397 | -24.3986
24.46 | -9.6422 | -7.5338 | -10.8648 | -23.9398
24.5875 | -9.6917 | -7.6954 | -10.9825 | -23.1093
24.715 | -9.7653 | -7.8736 | -11.1006 | -21.6548
24.8425 | -9.8573 | -8.0469 | -11.2742 | -19.3466
24.97 | -9.9685 | -8.2341 | -11.4681 | -16.4565
25.0975 | -10.0843 | -8.428 | -11.669 | -14.6385
25.225 | -10.2072 | -8.6404 | -11.8803 | -13.1496
25.3525 | -10.2975 | -8.8618 | -12.095 | -11.9492
25.48 | -10.4049 | -9.0799 | -12.3337 | -11.0276
25.6075 | -10.4874 | -9.3031 | -12.5951 | -10.2267
25.735 | -10.6076 | -9.5163 | -12.8638 | -9.5113
25.8625 | -10.6968 | -9.7064 | -13.1381 | -8.9351
25.99 | -10.8442 | -9.8633 | -13.4583 | -8.4237
26.1175 | -10.9535 | -9.9939 | -13.768 | -8.0147
26.245 | -11.0967 | -10.0992 | -14.0589 | -7.6631
26.3725 | -11.1941 | -10.1763 | -14.3038 | -7.3925
26.5 | -11.2919 | -10.2119 | -14.5039 | -7.1954

## CL vs. LO
RF (GHz) |  |  |  | A Configuration | [5 empty] | B Configuration
| +22 dBm | +17 dBm | +15.5 dBm | +14.5 dBm | +13.5 dBm | +12.5 dBm | +22 dBm | +17 dBm | +15.5 dBm | +14.5 dBm | +13.5 dBm | +12.5 dBm
1 | -68.2156 | -70.228 | -76.501 | -72.9176 | -74.7873 | -69.8363 | -23.325 | -60.9823 | -70.9167 | -73.9869 | -78.9111 | -72.2788
1.1275 | -69.3521 | -70.6572 | -75.1054 | -73.1609 | -75.7636 | -73.1959 | -22.5179 | -56.3718 | -71.1333 | -74.0508 | -76.6993 | -71.6357
1.255 | -66.1872 | -77.5854 | -75.7212 | -72.6317 | -76.8301 | -73.3598 | -21.4856 | -48.4749 | -65.841 | -72.7173 | -75.075 | -74.0463
1.3825 | -59.2361 | -78.7347 | -74.4343 | -71.7378 | -75.7629 | -73.5896 | -20.3104 | -39.2313 | -57.3415 | -66.3878 | -69.9272 | -73.4468
1.51 | -52.1325 | -79.5112 | -73.5838 | -69.392 | -76.259 | -73.1863 | -19.3313 | -34.0756 | -51.8235 | -63.1222 | -69.8002 | -73.5561
1.6375 | -42.5674 | -78.7316 | -71.4764 | -69.4005 | -80.0422 | -72.6414 | -18.5021 | -28.9277 | -45.5092 | -58.2972 | -68.3169 | -73.5868
1.765 | -31.9681 | -77.2277 | -71.9656 | -69.66 | -77.7916 | -69.4441 | -17.8513 | -23.7993 | -35.453 | -49.6567 | -62.736 | -71.9937
1.8925 | -25.0875 | -70.4676 | -69.1914 | -69.3512 | -75.2478 | -68.7735 | -17.3485 | -21.9349 | -30.9552 | -42.6663 | -57.0129 | -67.6507
2.02 | -22.1038 | -67.9525 | -69.7339 | -68.715 | -74.9185 | -68.7043 | -17.0157 | -21.3966 | -29.5042 | -40.4435 | -54.6854 | -65.4871
2.1475 | -17.3536 | -63.5481 | -68.8085 | -68.5686 | -73.6424 | -69.4456 | -16.6352 | -19.7777 | -25.0575 | -33.5423 | -47.5774 | -61.5724
2.275 | -15.1215 | -59.2989 | -68.1821 | -67.6602 | -68.5467 | -68.6247 | -16.294 | -18.7306 | -21.7999 | -27.4242 | -39.7108 | -55.7812
2.4025 | -14.5175 | -53.5342 | -65.3469 | -66.3215 | -68.6117 | -68.4654 | -16.1356 | -18.6405 | -21.2479 | -26.0233 | -37.0553 | -52.2795
2.53 | -14.3264 | -45.5562 | -62.9038 | -65.6411 | -68.0443 | -69.6885 | -15.8379 | -18.4582 | -20.5824 | -24.1977 | -33.1142 | -46.6718
2.6575 | -14.014 | -37.6848 | -56.3293 | -64.0788 | -66.7248 | -68.4283 | -15.4732 | -18.2645 | -20.1379 | -22.681 | -29.2644 | -41.4927
2.785 | -14.0293 | -30.4615 | -49.4195 | -59.8627 | -65.8386 | -66.4498 | -15.0624 | -17.9963 | -19.9487 | -21.9213 | -26.5592 | -36.4633
2.9125 | -14.5694 | -26.0606 | -41.6932 | -54.6636 | -64.3424 | -65.4798 | -14.6343 | -17.6909 | -20.1236 | -22.2056 | -26.2019 | -34.3547
3.04 | -14.1128 | -23.3843 | -35.4459 | -47.6723 | -59.329 | -63.7532 | -13.7531 | -17.0229 | -19.8907 | -22.2015 | -25.4583 | -31.5425
3.1675 | -13.5965 | -21.4666 | -29.8251 | -40.8097 | -54.178 | -60.3396 | -13.0197 | -16.6175 | -19.798 | -22.328 | -25.4209 | -30.3221
3.295 | -12.9464 | -20.296 | -26.1233 | -34.2644 | -47.1371 | -55.8617 | -12.1448 | -15.6785 | -19.0173 | -21.7556 | -24.8469 | -28.834
3.4225 | -12.332 | -19.3836 | -23.8952 | -29.4937 | -40.0222 | -52.031 | -11.326 | -15.0352 | -18.4912 | -21.3329 | -24.3497 | -27.7589
3.55 | -10.8892 | -17.7008 | -21.3885 | -24.8707 | -32.3519 | -44.1894 | -10.5462 | -14.4742 | -17.5821 | -20.1774 | -22.8546 | -25.5396
3.6775 | -10.1307 | -15.9986 | -19.1939 | -21.6823 | -26.7896 | -36.5772 | -10.1909 | -14.2072 | -16.8887 | -19.01 | -21.1626 | -23.2599
3.805 | -9.6467 | -14.3881 | -16.984 | -18.7103 | -21.748 | -29.0649 | -9.7358 | -13.3928 | -15.7177 | -17.4741 | -19.1712 | -20.7323
3.9325 | -9.3217 | -13.4131 | -15.4366 | -16.79 | -18.9948 | -24.0006 | -9.3607 | -12.8521 | -14.8283 | -16.2098 | -17.4277 | -18.5901
4.06 | -8.8602 | -12.2147 | -13.8003 | -14.8451 | -15.9582 | -18.2939 | -9.0654 | -12.1736 | -13.6665 | -14.6442 | -15.5261 | -16.5077
4.1875 | -8.7908 | -11.5848 | -12.8448 | -13.637 | -14.4723 | -16.1407 | -8.6753 | -11.2918 | -12.5582 | -13.3732 | -14.0949 | -15.0416
4.315 | -8.7745 | -11.0795 | -12.0007 | -12.5765 | -13.176 | -14.154 | -8.2137 | -10.219 | -11.3113 | -12.0569 | -12.763 | -13.725
4.4425 | -8.6861 | -10.6162 | -11.3207 | -11.7689 | -12.2508 | -12.9652 | -7.9015 | -9.4753 | -10.3598 | -11.0174 | -11.7297 | -12.7386
4.57 | -8.4991 | -10.082 | -10.6636 | -11.0538 | -11.4844 | -12.0854 | -7.6889 | -8.9985 | -9.69 | -10.227 | -10.888 | -11.8563
4.6975 | -8.3476 | -9.731 | -10.1919 | -10.5172 | -10.8992 | -11.4181 | -7.4776 | -8.5275 | -9.1174 | -9.6187 | -10.2822 | -11.2603
4.825 | -8.1689 | -9.3756 | -9.7832 | -10.0863 | -10.4459 | -10.9133 | -7.398 | -8.2413 | -8.7134 | -9.1536 | -9.7762 | -10.7002
4.9525 | -7.9532 | -9.0354 | -9.434 | -9.73 | -10.0794 | -10.5205 | -7.3737 | -8.1271 | -8.5011 | -8.8577 | -9.4069 | -10.2669
5.08 | -7.7351 | -8.7157 | -9.1151 | -9.408 | -9.7491 | -10.1607 | -7.3779 | -8.07 | -8.3779 | -8.6783 | -9.1595 | -9.9536
5.2075 | -7.583 | -8.4553 | -8.8386 | -9.1232 | -9.456 | -9.8465 | -7.4293 | -8.074 | -8.342 | -8.6129 | -9.0496 | -9.8012
5.335 | -7.4687 | -8.2602 | -8.6354 | -8.9145 | -9.2353 | -9.6064 | -7.4885 | -8.0831 | -8.3164 | -8.5481 | -8.9235 | -9.5714
5.4625 | -7.3517 | -8.0719 | -8.4602 | -8.7523 | -9.0858 | -9.4692 | -7.5596 | -8.1141 | -8.3391 | -8.5607 | -8.9125 | -9.5064
5.59 | -7.2581 | -7.8768 | -8.2474 | -8.5552 | -8.9264 | -9.3516 | -7.6667 | -8.1545 | -8.3626 | -8.5765 | -8.9147 | -9.4578
5.7175 | -7.1889 | -7.7165 | -8.0506 | -8.3485 | -8.7304 | -9.1952 | -7.7935 | -8.2256 | -8.4168 | -8.621 | -8.9329 | -9.4139
5.845 | -7.1218 | -7.5986 | -7.9133 | -8.2065 | -8.5987 | -9.1148 | -7.8782 | -8.2508 | -8.4391 | -8.6419 | -8.944 | -9.3959
5.9725 | -7.0578 | -7.4858 | -7.7865 | -8.0822 | -8.4979 | -9.0826 | -7.9489 | -8.2447 | -8.4385 | -8.6433 | -8.9468 | -9.3885
6.1 | -7.0025 | -7.398 | -7.6673 | -7.9542 | -8.3928 | -9.0465 | -7.9875 | -8.2241 | -8.4138 | -8.6167 | -8.9171 | -9.365
6.2275 | -6.9561 | -7.3342 | -7.5929 | -7.8679 | -8.3035 | -9.0162 | -8.0033 | -8.1879 | -8.3779 | -8.5844 | -8.8885 | -9.3497
6.355 | -6.9039 | -7.2864 | -7.5516 | -7.8331 | -8.2849 | -9.0527 | -7.9912 | -8.137 | -8.335 | -8.5522 | -8.8847 | -9.3812
6.4825 | -6.8623 | -7.2427 | -7.5128 | -7.8051 | -8.2803 | -9.0974 | -7.9592 | -8.0941 | -8.3061 | -8.5342 | -8.884 | -9.3994
6.61 | -6.824 | -7.1906 | -7.4683 | -7.7803 | -8.3074 | -9.2421 | -7.9223 | -8.0674 | -8.2893 | -8.5299 | -8.8886 | -9.4238
6.7375 | -6.7925 | -7.1349 | -7.4187 | -7.7369 | -8.2821 | -9.2566 | -7.892 | -8.0714 | -8.3083 | -8.5622 | -8.9354 | -9.4973
6.865 | -6.78 | -7.1299 | -7.4431 | -7.7981 | -8.4148 | -9.5396 | -7.8648 | -8.0879 | -8.3408 | -8.6016 | -8.9848 | -9.5614
6.9925 | -6.7795 | -7.1355 | -7.4783 | -7.8765 | -8.577 | -9.881 | -7.83 | -8.1027 | -8.3686 | -8.6347 | -9.016 | -9.5902
7.12 | -6.7893 | -7.1498 | -7.5144 | -7.942 | -8.7103 | -10.1794 | -7.8135 | -8.1354 | -8.4122 | -8.6881 | -9.09 | -9.6856
7.2475 | -6.8074 | -7.1784 | -7.5751 | -8.0437 | -8.9215 | -10.6392 | -7.81 | -8.1797 | -8.4642 | -8.7493 | -9.1652 | -9.7745
7.375 | -6.837 | -7.2198 | -7.6652 | -8.212 | -9.2718 | -11.3949 | -7.8154 | -8.2298 | -8.5244 | -8.8219 | -9.2635 | -9.8864
7.5025 | -6.88 | -7.2673 | -7.7355 | -8.3125 | -9.4451 | -11.7241 | -7.8397 | -8.2999 | -8.6074 | -8.9245 | -9.4018 | -10.0702
7.63 | -6.9346 | -7.3704 | -7.9086 | -8.5943 | -9.9533 | -12.6805 | -7.899 | -8.4083 | -8.7276 | -9.0603 | -9.5591 | -10.2623
7.7575 | -6.9925 | -7.4588 | -8.0632 | -8.857 | -10.4421 | -13.577 | -7.9736 | -8.5378 | -8.8733 | -9.2244 | -9.744 | -10.4884
7.885 | -7.0559 | -7.5673 | -8.227 | -9.1147 | -10.8707 | -14.295 | -8.0696 | -8.6877 | -9.0496 | -9.4324 | -9.9935 | -10.7908
8.0125 | -7.1173 | -7.7224 | -8.4554 | -9.447 | -11.3507 | -15.0476 | -8.1958 | -8.8541 | -9.2324 | -9.6311 | -10.2244 | -11.0655
8.14 | -7.1793 | -7.9579 | -8.856 | -10.0827 | -12.353 | -17.0352 | -8.3448 | -9.0718 | -9.473 | -9.8958 | -10.5243 | -11.3996
8.2675 | -7.2342 | -8.1 | -9.0423 | -10.3115 | -12.6174 | -17.2869 | -8.5121 | -9.316 | -9.7517 | -10.2138 | -10.8944 | -11.8244
8.395 | -7.2938 | -8.3125 | -9.3714 | -10.797 | -13.3217 | -18.5876 | -8.7133 | -9.5844 | -10.0675 | -10.5722 | -11.2791 | -12.2144
8.5225 | -7.3442 | -8.5275 | -9.7268 | -11.367 | -14.1624 | -20.3478 | -8.9046 | -9.8355 | -10.3629 | -10.8954 | -11.618 | -12.5574
8.65 | -7.3983 | -8.6781 | -9.9211 | -11.6388 | -14.469 | -20.7936 | -9.0757 | -10.0566 | -10.6386 | -11.214 | -11.9564 | -12.9078
8.7775 | -7.4435 | -8.7657 | -10.0598 | -11.8185 | -14.7114 | -21.4178 | -9.2216 | -10.2444 | -10.8638 | -11.4597 | -12.2075 | -13.176
8.905 | -7.4926 | -8.857 | -10.216 | -12.071 | -15.0368 | -22.3333 | -9.3244 | -10.3687 | -11.0083 | -11.6041 | -12.3526 | -13.3359
9.0325 | -7.5345 | -8.9413 | -10.44 | -12.3564 | -15.6493 | -24.0747 | -9.3656 | -10.4006 | -11.0343 | -11.6139 | -12.3796 | -13.4197
9.16 | -7.5854 | -8.9575 | -10.4793 | -12.2901 | -15.4215 | -23.7995 | -9.3887 | -10.3908 | -11.0063 | -11.5647 | -12.3335 | -13.4243
9.2875 | -7.6177 | -8.9934 | -10.6354 | -12.4313 | -15.7501 | -25.1768 | -9.3946 | -10.3785 | -10.9725 | -11.5115 | -12.2841 | -13.48
9.415 | -7.6376 | -9.0292 | -10.7267 | -12.4842 | -16.2105 | -26.4714 | -9.3718 | -10.2986 | -10.8649 | -11.3785 | -12.1338 | -13.3549
9.5425 | -7.6463 | -9.0462 | -10.7876 | -12.4597 | -16.277 | -27.1043 | -9.3205 | -10.1836 | -10.7366 | -11.2384 | -12.0015 | -13.3337
9.67 | -7.6543 | -9.0664 | -10.6964 | -12.2449 | -15.9334 | -26.6444 | -9.2827 | -10.097 | -10.6395 | -11.1281 | -11.8725 | -13.2567
9.7975 | -7.654 | -9.1497 | -10.7273 | -12.277 | -16.4107 | -28.0155 | -9.2372 | -10.0473 | -10.5907 | -11.071 | -11.8086 | -13.2494
9.925 | -7.6592 | -9.1736 | -10.6213 | -12.0961 | -16.1765 | -27.6437 | -9.1738 | -9.9959 | -10.5367 | -11.0136 | -11.7919 | -13.3369
10.0525 | -7.679 | -9.2045 | -10.5164 | -11.9793 | -16.0751 | -27.5381 | -9.1151 | -9.9251 | -10.4739 | -10.9526 | -11.7506 | -13.3733
10.18 | -7.6972 | -9.1983 | -10.3825 | -11.7687 | -15.6942 | -26.5982 | -9.0484 | -9.8375 | -10.3953 | -10.8871 | -11.7477 | -13.5009
10.3075 | -7.7183 | -9.2032 | -10.3243 | -11.7899 | -16.0777 | -27.3298 | -8.9645 | -9.7256 | -10.2905 | -10.7957 | -11.7257 | -13.6013
10.435 | -7.7336 | -9.1189 | -10.1414 | -11.4594 | -15.1653 | -25.0601 | -8.8731 | -9.6096 | -10.1845 | -10.7322 | -11.788 | -13.858
10.5625 | -7.7576 | -9.092 | -10.0845 | -11.4007 | -15.0732 | -24.8485 | -8.7843 | -9.4761 | -10.0577 | -10.6495 | -11.7765 | -13.9372
10.69 | -7.7608 | -8.9821 | -9.9283 | -11.1147 | -14.1901 | -22.8225 | -8.6892 | -9.3384 | -9.9233 | -10.545 | -11.7469 | -14.0091
10.8175 | -7.7576 | -8.9171 | -9.8698 | -11.0632 | -14.075 | -22.3591 | -8.6082 | -9.2395 | -9.854 | -10.5546 | -11.8978 | -14.3654
10.945 | -7.7416 | -8.7916 | -9.7332 | -10.7967 | -13.1179 | -20.0858 | -8.5218 | -9.1304 | -9.7593 | -10.5067 | -11.9185 | -14.4866
11.0725 | -7.7216 | -8.6752 | -9.6382 | -10.739 | -13.0389 | -19.7131 | -8.4465 | -9.0442 | -9.7079 | -10.5174 | -12.0095 | -14.7508
11.2 | -7.6943 | -8.5428 | -9.5019 | -10.5952 | -12.725 | -18.6076 | -8.3844 | -8.9525 | -9.6688 | -10.5579 | -12.1837 | -15.1857
11.3275 | -7.6701 | -8.451 | -9.4 | -10.4877 | -12.4955 | -17.7 | -8.3293 | -8.8723 | -9.6175 | -10.5882 | -12.3498 | -15.6084
11.455 | -7.6783 | -8.4332 | -9.3696 | -10.446 | -12.3835 | -17.1467 | -8.2978 | -8.8455 | -9.6707 | -10.7865 | -12.7984 | -16.5932
11.5825 | -7.7048 | -8.4391 | -9.3726 | -10.4533 | -12.3302 | -16.7141 | -8.2944 | -8.8392 | -9.6703 | -10.8107 | -12.844 | -16.6816
11.71 | -7.7438 | -8.4738 | -9.3821 | -10.4344 | -12.235 | -16.2255 | -8.303 | -8.8907 | -9.83 | -11.1255 | -13.4917 | -18.1554
11.8375 | -7.8059 | -8.5692 | -9.4927 | -10.5387 | -12.2829 | -15.9855 | -8.3226 | -8.9062 | -9.8373 | -11.1476 | -13.5302 | -18.3187
11.965 | -7.9289 | -8.6989 | -9.6304 | -10.6781 | -12.3824 | -15.8601 | -8.3794 | -9.0024 | -10.003 | -11.4033 | -13.9644 | -19.2828
12.0925 | -8.0517 | -8.8178 | -9.7744 | -10.8854 | -12.6647 | -16.2774 | -8.4454 | -9.1061 | -10.1054 | -11.4325 | -13.8267 | -18.9976
12.22 | -8.1552 | -8.9144 | -9.851 | -10.9697 | -12.7355 | -16.1354 | -8.5111 | -9.2885 | -10.3903 | -11.8074 | -14.372 | -20.1462
12.3475 | -8.2866 | -9.0747 | -10.0538 | -11.2411 | -13.1366 | -16.8858 | -8.5958 | -9.4341 | -10.5197 | -11.874 | -14.2906 | -19.9551
12.475 | -8.3822 | -9.1661 | -10.1644 | -11.4098 | -13.4338 | -17.5742 | -8.6943 | -9.6113 | -10.7417 | -12.091 | -14.5014 | -20.3406
12.6025 | -8.4141 | -9.2133 | -10.2439 | -11.5674 | -13.7956 | -18.6814 | -8.756 | -9.7613 | -10.8948 | -12.2026 | -14.5055 | -20.3311
12.73 | -8.3921 | -9.2038 | -10.229 | -11.5559 | -13.8659 | -19.1502 | -8.7843 | -9.8797 | -11.0023 | -12.3077 | -14.6398 | -20.7884
12.8575 | -8.3652 | -9.2149 | -10.3155 | -11.7511 | -14.4584 | -21.0839 | -8.7958 | -9.9031 | -10.966 | -12.2074 | -14.4065 | -20.2259
12.985 | -8.2877 | -9.1403 | -10.2321 | -11.6527 | -14.3729 | -21.3066 | -8.7928 | -9.8925 | -10.9063 | -12.0963 | -14.1654 | -19.7903
13.1125 | -8.2261 | -9.1317 | -10.2535 | -11.7196 | -14.6752 | -22.5131 | -8.7581 | -9.8714 | -10.8673 | -12.057 | -14.0913 | -19.7507
13.24 | -8.1583 | -9.0992 | -10.2187 | -11.6736 | -14.6752 | -22.8665 | -8.7334 | -9.8567 | -10.8697 | -12.0872 | -14.2408 | -20.4348
13.3675 | -8.1188 | -9.1018 | -10.2711 | -11.7964 | -15.0899 | -24.2703 | -8.7227 | -9.7855 | -10.7557 | -11.9116 | -13.8889 | -19.471
13.495 | -8.0835 | -9.1032 | -10.2502 | -11.7104 | -14.8911 | -23.9411 | -8.7276 | -9.8043 | -10.8187 | -12.0209 | -14.1451 | -20.4384
13.6225 | -8.0625 | -9.1418 | -10.3177 | -11.8054 | -15.1483 | -24.8275 | -8.7251 | -9.8032 | -10.8447 | -12.071 | -14.3052 | -21.0384
13.75 | -8.0452 | -9.1912 | -10.4148 | -11.9888 | -15.8575 | -26.2774 | -8.7494 | -9.8569 | -10.9597 | -12.235 | -14.7384 | -22.3865
13.8775 | -8.0335 | -9.2371 | -10.5166 | -12.1846 | -16.5809 | -27.9308 | -8.7929 | -9.8564 | -10.9224 | -12.1437 | -14.4604 | -21.4741
14.005 | -8.0399 | -9.286 | -10.541 | -12.1554 | -16.4201 | -27.5064 | -8.8629 | -9.9164 | -11.0229 | -12.2985 | -14.736 | -22.2897
14.1325 | -8.0534 | -9.4481 | -10.8533 | -12.9447 | -18.6888 | -31.0715 | -8.9428 | -10.0156 | -11.2538 | -12.7065 | -15.9627 | -24.9672
14.26 | -8.1209 | -9.5397 | -11.0134 | -13.2576 | -19.7262 | -33.1947 | -9.0481 | -10.005 | -11.1385 | -12.4945 | -15.5043 | -23.3207
14.3875 | -8.1923 | -9.5918 | -11.0617 | -13.3057 | -19.8048 | -33.3491 | -9.1767 | -10.0527 | -11.1827 | -12.6001 | -15.6365 | -23.5137
14.515 | -8.2314 | -9.7645 | -11.3725 | -14.0442 | -21.7201 | -35.9158 | -9.2652 | -10.1366 | -11.2711 | -12.7692 | -15.9474 | -24.1968
14.6425 | -8.2989 | -9.9134 | -11.73 | -14.9703 | -24.3987 | -40.1405 | -9.3813 | -10.2016 | -11.2426 | -12.6625 | -15.7588 | -23.341
14.77 | -8.3281 | -9.8476 | -11.5896 | -14.4926 | -23.1042 | -38.4901 | -9.4725 | -10.1876 | -11.0045 | -12.185 | -14.3646 | -19.7034
14.8975 | -8.2678 | -9.9375 | -11.8853 | -15.4021 | -25.2993 | -41.133 | -9.5389 | -10.3219 | -11.1566 | -12.3792 | -14.7316 | -20.6262
15.025 | -8.2062 | -9.9177 | -11.9874 | -15.8253 | -26.5912 | -42.8686 | -9.5854 | -10.362 | -11.0419 | -12.0278 | -13.9284 | -18.2349
15.1525 | -8.2005 | -9.7889 | -11.827 | -15.5258 | -25.9776 | -42.2372 | -9.6589 | -10.4067 | -10.995 | -11.8035 | -13.4777 | -16.9766
15.28 | -8.111 | -9.6429 | -11.6271 | -15.1441 | -25.1394 | -41.2809 | -9.6751 | -10.4388 | -10.9988 | -11.7282 | -13.2365 | -16.4996
15.4075 | -8.0605 | -9.5817 | -11.6427 | -15.4677 | -26.1798 | -42.7023 | -9.7141 | -10.4904 | -11.0223 | -11.6892 | -13.0791 | -16.287
15.535 | -8.0755 | -9.4187 | -11.3637 | -14.9152 | -25.1122 | -41.6186 | -9.7655 | -10.4743 | -10.9444 | -11.4854 | -12.5721 | -15.0762
15.6625 | -8.1092 | -9.2921 | -11.1328 | -14.5359 | -24.2951 | -40.5925 | -9.8321 | -10.4825 | -10.9416 | -11.4614 | -12.501 | -14.9519
15.79 | -8.0505 | -9.1622 | -10.9326 | -14.2911 | -23.8893 | -40.0566 | -9.8608 | -10.4736 | -10.9022 | -11.3718 | -12.2728 | -14.3818
15.9175 | -8.0317 | -9.091 | -10.8369 | -14.2919 | -24.0699 | -40.155 | -9.8703 | -10.4591 | -10.8751 | -11.3378 | -12.2194 | -14.2797
16.045 | -8.0724 | -8.9844 | -10.6328 | -13.9489 | -23.5232 | -39.3639 | -9.9152 | -10.4078 | -10.7933 | -11.2056 | -11.9406 | -13.5544
16.1725 | -8.0059 | -8.8278 | -10.3564 | -13.4156 | -22.5397 | -38.0763 | -9.8943 | -10.3504 | -10.7301 | -11.1382 | -11.8563 | -13.4319
16.3 | -7.8695 | -8.7466 | -10.2103 | -13.1337 | -22.0637 | -37.4949 | -9.7872 | -10.2753 | -10.6443 | -11.0414 | -11.7064 | -13.1041
16.4275 | -7.8415 | -8.6963 | -10.12 | -13.004 | -22.0276 | -37.5161 | -9.7187 | -10.1866 | -10.5497 | -10.942 | -11.5857 | -12.9126
16.555 | -7.7999 | -8.5652 | -9.8654 | -12.4077 | -20.7837 | -36.0773 | -9.6632 | -10.0779 | -10.4233 | -10.7892 | -11.371 | -12.5073
16.6825 | -7.6559 | -8.4543 | -9.639 | -11.8649 | -19.324 | -34.0733 | -9.5071 | -9.9524 | -10.2938 | -10.6549 | -11.231 | -12.3402
16.81 | -7.597 | -8.4512 | -9.6123 | -11.8001 | -19.2719 | -34.1455 | -9.4017 | -9.8677 | -10.197 | -10.545 | -11.1001 | -12.1362
16.9375 | -7.5777 | -8.3583 | -9.4454 | -11.4498 | -18.2695 | -32.6166 | -9.3426 | -9.757 | -10.0671 | -10.3942 | -10.907 | -11.8281
17.065 | -7.5351 | -8.2548 | -9.2252 | -10.865 | -16.4266 | -29.6269 | -9.2658 | -9.6555 | -9.9545 | -10.2648 | -10.7468 | -11.5844
17.1925 | -7.4803 | -8.212 | -9.1124 | -10.5632 | -15.3352 | -27.6615 | -9.1786 | -9.5685 | -9.8582 | -10.1621 | -10.6316 | -11.4417
17.32 | -7.4534 | -8.1569 | -9.0293 | -10.4112 | -14.9017 | -26.8616 | -9.1272 | -9.4828 | -9.7595 | -10.0422 | -10.4766 | -11.1979
17.4475 | -7.4622 | -8.0607 | -8.8147 | -9.8764 | -13.0263 | -22.9804 | -9.0724 | -9.3771 | -9.6428 | -9.9112 | -10.3175 | -10.9798
17.575 | -7.4631 | -8.0249 | -8.726 | -9.6717 | -12.2884 | -21.0104 | -9.0124 | -9.296 | -9.5562 | -9.8135 | -10.2109 | -10.8578
17.7025 | -7.4296 | -7.9832 | -8.6393 | -9.4991 | -11.7259 | -19.2789 | -8.9397 | -9.2175 | -9.4664 | -9.7153 | -10.0902 | -10.6953
17.83 | -7.4604 | -7.9615 | -8.5801 | -9.3566 | -11.2128 | -17.7016 | -8.8988 | -9.1412 | -9.3837 | -9.6202 | -9.974 | -10.534
17.9575 | -7.4903 | -7.929 | -8.4815 | -9.1294 | -10.4613 | -15.1045 | -8.8509 | -9.0681 | -9.3011 | -9.5321 | -9.8775 | -10.4211
18.085 | -7.4602 | -7.9207 | -8.4575 | -9.0713 | -10.2764 | -14.3555 | -8.7945 | -9.0141 | -9.2392 | -9.46 | -9.7933 | -10.3174
18.2125 | -7.4644 | -7.9349 | -8.4358 | -8.985 | -9.9863 | -13.1048 | -8.7494 | -8.9594 | -9.178 | -9.3907 | -9.7121 | -10.2089
18.34 | -7.4885 | -7.9451 | -8.4131 | -8.9077 | -9.7649 | -12.1985 | -8.7123 | -8.9035 | -9.1147 | -9.3167 | -9.6294 | -10.1033
18.4675 | -7.4787 | -7.9546 | -8.3877 | -8.8221 | -9.4995 | -11.0507 | -8.6816 | -8.8708 | -9.076 | -9.2745 | -9.5849 | -10.0615
18.595 | -7.501 | -8.0122 | -8.431 | -8.8396 | -9.4604 | -10.8264 | -8.6795 | -8.8546 | -9.0521 | -9.2469 | -9.5468 | -10.0078
18.7225 | -7.5421 | -8.0482 | -8.4352 | -8.7974 | -9.3106 | -10.2252 | -8.6973 | -8.8556 | -9.0494 | -9.2487 | -9.5528 | -10.02
18.85 | -7.5642 | -8.078 | -8.4506 | -8.7921 | -9.2567 | -10.0278 | -8.712 | -8.8582 | -9.0465 | -9.2453 | -9.5515 | -10.0306
18.9775 | -7.6144 | -8.1364 | -8.4954 | -8.8152 | -9.2327 | -9.8793 | -8.7508 | -8.8827 | -9.0696 | -9.2751 | -9.5883 | -10.0898
19.105 | -7.6739 | -8.1893 | -8.5332 | -8.8369 | -9.218 | -9.7777 | -8.7809 | -8.9058 | -9.0922 | -9.2975 | -9.6178 | -10.1305
19.2325 | -7.7217 | -8.2326 | -8.5664 | -8.8558 | -9.2067 | -9.7018 | -8.8134 | -8.9387 | -9.1339 | -9.3506 | -9.6937 | -10.2501
19.36 | -7.8069 | -8.3159 | -8.6424 | -8.9209 | -9.2527 | -9.7069 | -8.8465 | -8.9641 | -9.163 | -9.383 | -9.7374 | -10.3188
19.4875 | -7.9135 | -8.4138 | -8.7375 | -9.0054 | -9.327 | -9.7541 | -8.9027 | -9.0218 | -9.2267 | -9.4573 | -9.8292 | -10.4409
19.615 | -8.001 | -8.5083 | -8.8296 | -9.0925 | -9.4028 | -9.8067 | -8.9444 | -9.0649 | -9.2781 | -9.5171 | -9.8995 | -10.5342
19.7425 | -8.1253 | -8.6296 | -8.9479 | -9.2039 | -9.5024 | -9.8864 | -8.9963 | -9.1193 | -9.3407 | -9.594 | -9.9933 | -10.6607
19.87 | -8.2652 | -8.7753 | -9.0929 | -9.3455 | -9.6383 | -10.0139 | -9.058 | -9.1817 | -9.4073 | -9.6627 | -10.0645 | -10.7439
19.9975 | -8.3689 | -8.8914 | -9.2043 | -9.4508 | -9.734 | -10.0984 | -9.1015 | -9.2268 | -9.4585 | -9.7189 | -10.1283 | -10.824
20.125 | -8.4927 | -9.0159 | -9.3258 | -9.567 | -9.8424 | -10.1966 | -9.152 | -9.2724 | -9.5091 | -9.7807 | -10.1991 | -10.9306
20.2525 | -8.6295 | -9.1448 | -9.449 | -9.6858 | -9.955 | -10.3002 | -9.2116 | -9.3314 | -9.5709 | -9.8496 | -10.2889 | -11.0696
20.38 | -8.7051 | -9.2243 | -9.5235 | -9.755 | -10.0169 | -10.3489 | -9.2734 | -9.3848 | -9.633 | -9.9191 | -10.3825 | -11.2247
20.5075 | -8.7763 | -9.3073 | -9.6033 | -9.8279 | -10.0847 | -10.4112 | -9.3165 | -9.4129 | -9.6668 | -9.9679 | -10.464 | -11.4027
20.635 | -8.8508 | -9.3864 | -9.6815 | -9.9019 | -10.1549 | -10.4745 | -9.3874 | -9.4802 | -9.7494 | -10.0767 | -10.6204 | -11.741
20.7625 | -8.8812 | -9.4416 | -9.7307 | -9.9473 | -10.1907 | -10.5086 | -9.4258 | -9.5133 | -9.7962 | -10.1484 | -10.7494 | -12.1425
20.89 | -8.9486 | -9.5344 | -9.8252 | -10.0368 | -10.276 | -10.5998 | -9.5035 | -9.5795 | -9.883 | -10.2654 | -10.9332 | -12.686
21.0175 | -9.0539 | -9.6636 | -9.9529 | -10.1609 | -10.3958 | -10.7328 | -9.5509 | -9.6291 | -9.9623 | -10.3888 | -11.1796 | -13.6275
21.145 | -9.1118 | -9.7339 | -10.0187 | -10.2249 | -10.4546 | -10.8046 | -9.5879 | -9.6767 | -10.0449 | -10.5204 | -11.4961 | -14.9429
21.2725 | -9.1735 | -9.8054 | -10.0858 | -10.2884 | -10.518 | -10.8793 | -9.602 | -9.6903 | -10.0843 | -10.5977 | -11.7188 | -15.9298
21.4 | -9.2627 | -9.8856 | -10.1632 | -10.3621 | -10.5919 | -10.9644 | -9.6288 | -9.736 | -10.1666 | -10.7197 | -12.0312 | -17.1338
21.5275 | -9.2944 | -9.9033 | -10.1695 | -10.3604 | -10.5861 | -10.9651 | -9.5918 | -9.7091 | -10.158 | -10.7306 | -12.1617 | -17.8571
21.655 | -9.3311 | -9.9226 | -10.1728 | -10.355 | -10.5793 | -10.964 | -9.5915 | -9.719 | -10.1791 | -10.7732 | -12.3425 | -18.5613
21.7825 | -9.4077 | -9.9637 | -10.1965 | -10.3709 | -10.5957 | -10.9903 | -9.5958 | -9.7272 | -10.1856 | -10.7807 | -12.3813 | -18.6651
21.91 | -9.4745 | -9.9873 | -10.2032 | -10.3683 | -10.5892 | -10.9933 | -9.5993 | -9.7351 | -10.185 | -10.7642 | -12.3311 | -18.4912
22.0375 | -9.5207 | -10.0021 | -10.2049 | -10.3671 | -10.5962 | -11.0312 | -9.6089 | -9.7398 | -10.177 | -10.7413 | -12.2734 | -18.2813
22.165 | -9.5658 | -10.026 | -10.2196 | -10.3814 | -10.6257 | -11.0942 | -9.6524 | -9.7817 | -10.2253 | -10.8203 | -12.5578 | -19.1385
22.2925 | -9.5744 | -10.0144 | -10.2011 | -10.3637 | -10.6149 | -11.0965 | -9.6722 | -9.7792 | -10.1978 | -10.756 | -12.3501 | -18.3885
22.42 | -9.5782 | -10.0001 | -10.1877 | -10.3585 | -10.631 | -11.1606 | -9.7028 | -9.8003 | -10.2102 | -10.7705 | -12.3793 | -18.418
22.5475 | -9.5785 | -9.9941 | -10.1827 | -10.366 | -10.6601 | -11.2397 | -9.7631 | -9.8374 | -10.2584 | -10.8639 | -12.7339 | -19.5219
22.675 | -9.5644 | -9.9745 | -10.1632 | -10.357 | -10.6673 | -11.2808 | -9.837 | -9.8568 | -10.2774 | -10.9067 | -12.9028 | -20.036
22.8025 | -9.5597 | -9.9596 | -10.1512 | -10.3549 | -10.6805 | -11.3227 | -9.918 | -9.8722 | -10.2761 | -10.8772 | -12.711 | -19.3767
22.93 | -9.5552 | -9.9381 | -10.1419 | -10.3635 | -10.73 | -11.4673 | -9.9812 | -9.9205 | -10.3731 | -11.1184 | -13.6427 | -21.8502
23.0575 | -9.5642 | -9.9424 | -10.1469 | -10.3721 | -10.7523 | -11.5017 | -10.0695 | -9.9593 | -10.4176 | -11.1784 | -13.8256 | -22.3682
23.185 | -9.5554 | -9.9189 | -10.1284 | -10.3656 | -10.7709 | -11.5695 | -10.1479 | -9.9866 | -10.4523 | -11.2424 | -14.0602 | -22.9866
23.3125 | -9.5539 | -9.9048 | -10.1234 | -10.3682 | -10.7932 | -11.6308 | -10.2371 | -10.0359 | -10.5117 | -11.3303 | -14.3377 | -23.7298
23.44 | -9.537 | -9.874 | -10.101 | -10.359 | -10.8172 | -11.721 | -10.2534 | -10.0813 | -10.6004 | -11.5681 | -15.3548 | -26.2351
23.5675 | -9.5268 | -9.8516 | -10.0778 | -10.338 | -10.792 | -11.6635 | -10.3227 | -10.1091 | -10.6024 | -11.5028 | -15.0607 | -25.7315
23.695 | -9.4964 | -9.8027 | -10.0383 | -10.3131 | -10.7852 | -11.6873 | -10.4101 | -10.1497 | -10.6369 | -11.5348 | -15.1009 | -25.8658
23.8225 | -9.4981 | -9.7985 | -10.041 | -10.3193 | -10.7899 | -11.6797 | -10.4501 | -10.2068 | -10.6863 | -11.5772 | -15.17 | -26.1095
23.95 | -9.4951 | -9.7801 | -10.0243 | -10.3059 | -10.7805 | -11.6545 | -10.4788 | -10.2566 | -10.7213 | -11.5954 | -15.1377 | -26.0887
24.0775 | -9.512 | -9.7839 | -10.0263 | -10.3081 | -10.7629 | -11.5855 | -10.651 | -10.3427 | -10.7521 | -11.4702 | -14.2295 | -23.7483
24.205 | -9.5307 | -9.7903 | -10.0391 | -10.3243 | -10.7852 | -11.6237 | -10.7772 | -10.4248 | -10.8138 | -11.4895 | -14.0371 | -23.0145
24.3325 | -9.5642 | -9.814 | -10.0657 | -10.3541 | -10.8127 | -11.6354 | -10.8144 | -10.5192 | -10.8933 | -11.5461 | -14.0046 | -22.5759
24.46 | -9.5896 | -9.8323 | -10.0905 | -10.3842 | -10.8504 | -11.6801 | -10.9233 | -10.64 | -10.9903 | -11.5847 | -13.7382 | -21.2662
24.5875 | -9.6393 | -9.8859 | -10.155 | -10.4588 | -10.9362 | -11.7933 | -11.0665 | -10.7618 | -11.0832 | -11.6154 | -13.4722 | -19.7622
24.715 | -9.7283 | -9.9766 | -10.2591 | -10.5751 | -11.0806 | -11.9832 | -11.1447 | -10.9009 | -11.2186 | -11.7563 | -13.6785 | -19.4692
24.8425 | -9.8237 | -10.0773 | -10.3903 | -10.7198 | -11.2581 | -12.2292 | -11.2452 | -11.0928 | -11.3946 | -11.927 | -13.7714 | -18.5995
24.97 | -9.9363 | -10.2143 | -10.5588 | -10.9156 | -11.5126 | -12.6441 | -11.4342 | -11.2742 | -11.5523 | -12.0453 | -13.7479 | -17.6313
25.0975 | -10.0575 | -10.3579 | -10.7312 | -11.1154 | -11.7687 | -13.0554 | -11.637 | -11.4698 | -11.7276 | -12.2399 | -13.8731 | -16.9029
25.225 | -10.1901 | -10.5195 | -10.9299 | -11.3647 | -12.1286 | -13.7379 | -11.8057 | -11.6863 | -11.9305 | -12.4657 | -14.1021 | -16.6469
25.3525 | -10.2774 | -10.6391 | -11.0942 | -11.5939 | -12.5025 | -14.5423 | -11.9849 | -11.8914 | -12.1229 | -12.7138 | -14.2182 | -16.2486
25.48 | -10.3744 | -10.8119 | -11.3329 | -11.9546 | -13.1988 | -16.2095 | -12.2342 | -12.1127 | -12.3129 | -12.9119 | -14.2379 | -15.7671
25.6075 | -10.453 | -10.948 | -11.5335 | -12.2776 | -13.8955 | -17.9346 | -12.5005 | -12.3742 | -12.5808 | -13.2826 | -14.5491 | -15.8946
25.735 | -10.5643 | -11.1316 | -11.8202 | -12.7892 | -15.0862 | -20.7821 | -12.7651 | -12.6321 | -12.8506 | -13.6037 | -14.7103 | -15.8559
25.8625 | -10.6398 | -11.3039 | -12.1378 | -13.4546 | -16.7964 | -24.4702 | -13.0499 | -12.9081 | -13.1817 | -14.0011 | -14.9687 | -16.0081
25.99 | -10.7674 | -11.5609 | -12.6346 | -14.5879 | -19.6072 | -28.9709 | -13.3695 | -13.2109 | -13.5144 | -14.3272 | -15.1564 | -16.0406
26.1175 | -10.8783 | -11.7424 | -12.9775 | -15.3776 | -21.6141 | -32.1846 | -13.6657 | -13.5127 | -13.9337 | -14.7672 | -15.4965 | -16.3555
26.245 | -11.0265 | -12.08 | -13.9667 | -17.8733 | -25.9233 | -35.5539 | -13.9463 | -13.8151 | -14.3249 | -15.1115 | -15.7501 | -16.549
26.3725 | -11.1297 | -12.3563 | -14.8535 | -20.164 | -29.7212 | -37.7733 | -14.1755 | -14.0671 | -14.6562 | -15.3904 | -15.964 | -16.7354
26.5 | -11.2351 | -12.5994 | -15.6505 | -22.2157 | -32.8532 | -38.8246 | -14.3581 | -14.2647 | -14.897 | -15.5856 | -16.1221 | -16.8654

## LORL LRI RI
LO - GHz | LO RL A | LO-RF A | LO-IF A | RF-IF A | LO RL B | LO-RF B | LO-IF B | RF-IF B |  | LO - GHz | A LO-RF Amp | A LO-IF Ampl | B LO-RF Amp | B LO-IF Amp
Average => | -9.6768 | -52.7791 | -28.2976 | -35.6816 | -9.1269 | -50.2179 | -35.6777 | -28.4166 | Amplified | Average => | -49.1774 | -28.9823 | -45.4112 | -35.1757
Data to right
1 | -4.2087 | -66.3161 | -37.074 | -44.1 | -4.8018 | -66.1213 | -38.9914 | -37.1331 | used on | 1 | -65.5598 | -37.2618 | -55.8303 | -34.8279
1.1275 | -4.1081 | -66.1348 | -36.669 | -43.7741 | -4.8856 | -65.0598 | -38.0906 | -36.6992 | data sheet | 1.1275 | -64.6757 | -37.4674 | -55.3686 | -35.5536
1.255 | -3.9878 | -65.7077 | -36.3252 | -43.4884 | -4.949 | -63.9073 | -37.2677 | -36.2783 |  | 1.255 | -63.7094 | -37.7734 | -54.9154 | -36.2568
1.3825 | -3.8292 | -65.3031 | -35.9324 | -43.0548 | -4.9654 | -62.885 | -36.3855 | -35.7241 |  | 1.3825 | -62.7234 | -38.0856 | -54.4857 | -36.8305
1.51 | -3.6073 | -65.5308 | -35.4286 | -42.6372 | -4.7954 | -62.1857 | -36.3371 | -35.172 |  | 1.51 | -61.8051 | -38.6519 | -54.1296 | -37.4314
1.6375 | -3.3327 | -65.7644 | -35.087 | -42.5076 | -4.3837 | -62.2847 | -37.8995 | -34.8242 |  | 1.6375 | -61.0265 | -39.2678 | -53.918 | -37.9698
1.765 | -3.0858 | -65.9143 | -34.7517 | -41.195 | -3.8348 | -63.0459 | -39.5943 | -34.5063 |  | 1.765 | -60.39 | -39.5231 | -53.7689 | -38.4709
1.8925 | -2.8282 | -66.2251 | -34.3573 | -40.349 | -3.2734 | -64.0545 | -41.391 | -34.1852 |  | 1.8925 | -60.0647 | -39.5031 | -53.6759 | -38.9162
2.02 | -2.584 | -66.2791 | -34.0419 | -39.9531 | -2.7411 | -65.2054 | -43.6481 | -34.0191 |  | 2.02 | -59.7821 | -39.3693 | -53.561 | -39.2332
2.1475 | -2.4183 | -65.772 | -33.9652 | -39.7426 | -2.3953 | -66.0981 | -44.5956 | -33.986 |  | 2.1475 | -59.5658 | -39.1668 | -53.4762 | -39.3014
2.275 | -2.3224 | -64.5943 | -33.8518 | -38.7579 | -2.3147 | -65.9292 | -43.5611 | -33.8416 |  | 2.275 | -59.3471 | -38.9628 | -53.4252 | -39.2243
2.4025 | -2.2718 | -63.0823 | -33.8906 | -39.2947 | -2.512 | -64.3523 | -42.1447 | -33.7864 |  | 2.4025 | -59.2641 | -38.6978 | -53.2947 | -39.0523
2.53 | -2.2672 | -61.622 | -34.1199 | -39.2073 | -2.7921 | -62.3643 | -40.3914 | -33.6725 |  | 2.53 | -59.2179 | -38.475 | -53.1295 | -38.9523
2.6575 | -2.3752 | -60.2585 | -35.1219 | -37.7325 | -3.2951 | -60.0664 | -38.0964 | -33.5896 |  | 2.6575 | -59.2178 | -38.2536 | -52.9912 | -38.9088
2.785 | -2.5057 | -59.1465 | -35.944 | -36.4785 | -3.9796 | -58.7977 | -36.2921 | -33.2639 |  | 2.785 | -59.037 | -38.0161 | -52.841 | -38.7942
2.9125 | -2.6245 | -58.5841 | -36.4294 | -35.0061 | -4.5299 | -58.5653 | -35.201 | -32.873 |  | 2.9125 | -58.8386 | -37.7799 | -52.6256 | -38.6616
3.04 | -2.983 | -58.8454 | -36.9186 | -35.0802 | -5.4005 | -58.9286 | -35.8653 | -33.0789 |  | 3.04 | -58.6611 | -37.5206 | -52.3667 | -38.4141
3.1675 | -3.2561 | -59.19 | -36.9499 | -34.3407 | -6.2363 | -59.3243 | -35.7633 | -33.2193 |  | 3.1675 | -58.4922 | -37.1867 | -52.0864 | -38.1076
3.295 | -3.4844 | -60.307 | -36.0742 | -34.705 | -6.8934 | -59.8955 | -35.5167 | -33.3379 |  | 3.295 | -58.0052 | -36.7994 | -51.7511 | -37.8523
3.4225 | -3.7639 | -62.0595 | -35.2417 | -34.3702 | -7.4586 | -59.0898 | -34.8048 | -33.4953 |  | 3.4225 | -57.3647 | -36.381 | -51.4384 | -37.5257
3.55 | -4.0043 | -63.549 | -34.4015 | -34.443 | -8.0251 | -57.7678 | -34.2871 | -33.5877 |  | 3.55 | -56.8367 | -35.8908 | -51.2686 | -36.9131
3.6775 | -4.1494 | -64.7804 | -33.4641 | -32.7161 | -8.1307 | -56.9223 | -32.2782 | -32.9316 |  | 3.6775 | -56.6191 | -35.242 | -51.1633 | -35.9624
3.805 | -4.4093 | -65.9322 | -32.5672 | -31.6826 | -8.2926 | -55.7382 | -30.8929 | -32.1568 |  | 3.805 | -56.5711 | -34.3881 | -51.0767 | -34.7928
3.9325 | -4.6848 | -65.8057 | -31.5633 | -30.5056 | -8.3901 | -54.2391 | -29.5891 | -31.2565 |  | 3.9325 | -56.5804 | -33.4078 | -51.0441 | -33.6637
4.06 | -5.0937 | -64.6062 | -30.7249 | -29.5585 | -8.47 | -52.7432 | -28.739 | -30.3395 |  | 4.06 | -56.5933 | -32.3933 | -51.079 | -32.6659
4.1875 | -5.5443 | -63.3222 | -29.9311 | -29.0482 | -8.5671 | -51.4755 | -28.2297 | -29.5189 |  | 4.1875 | -56.5719 | -31.099 | -51.0908 | -31.6474
4.315 | -5.8624 | -61.38 | -29.0921 | -28.2845 | -8.5968 | -49.9857 | -27.6241 | -28.6904 |  | 4.315 | -56.4063 | -29.7313 | -50.9478 | -30.6975
4.4425 | -6.1635 | -59.0816 | -28.4167 | -27.5583 | -8.6273 | -48.6078 | -27.2576 | -27.822 |  | 4.4425 | -56.097 | -28.7045 | -50.6696 | -29.774
4.57 | -6.4253 | -56.6731 | -27.7287 | -27.0959 | -8.7119 | -47.1537 | -27.0545 | -27.0476 |  | 4.57 | -55.4742 | -27.9448 | -50.2685 | -28.9032
4.6975 | -6.5131 | -54.7729 | -27.062 | -26.7307 | -8.7251 | -46.1992 | -26.8579 | -26.3757 |  | 4.6975 | -54.8106 | -27.2715 | -49.8151 | -28.1475
4.825 | -6.6205 | -53.1999 | -26.5116 | -26.3165 | -8.8584 | -45.2423 | -26.6788 | -25.6743 |  | 4.825 | -54.1179 | -26.6601 | -49.3038 | -27.6449
4.9525 | -6.7715 | -51.5025 | -25.8528 | -26.1097 | -9.0882 | -44.2249 | -26.4888 | -24.9723 |  | 4.9525 | -53.4814 | -26.0492 | -48.6901 | -27.287
5.08 | -6.8975 | -50.0918 | -25.2031 | -26.2359 | -9.3048 | -43.6658 | -26.4167 | -24.6091 |  | 5.08 | -52.7221 | -25.4876 | -48.1449 | -27.0126
5.2075 | -7.0559 | -49.1275 | -24.6123 | -26.2296 | -9.5576 | -43.1627 | -26.3522 | -24.1021 |  | 5.2075 | -52.0162 | -24.9473 | -47.7832 | -26.7511
5.335 | -7.2151 | -48.3129 | -23.9269 | -26.2783 | -9.8711 | -42.8579 | -26.2293 | -23.6451 |  | 5.335 | -51.316 | -24.463 | -47.4044 | -26.4982
5.4625 | -7.3709 | -47.7606 | -23.3402 | -26.3306 | -10.1594 | -42.6963 | -26.1467 | -23.2124 |  | 5.4625 | -50.8127 | -24.0185 | -47.0237 | -26.2923
5.59 | -7.5365 | -47.4973 | -22.8444 | -26.4002 | -10.4652 | -42.8511 | -26.134 | -22.8656 |  | 5.59 | -50.4384 | -23.5949 | -46.7116 | -26.1474
5.7175 | -7.6952 | -47.2412 | -22.3809 | -26.2755 | -10.7502 | -43.2037 | -26.0682 | -22.3239 |  | 5.7175 | -50.1781 | -23.1942 | -46.514 | -26.0679
5.845 | -7.8753 | -47.0281 | -21.9882 | -26.0903 | -11.0887 | -43.8873 | -26.0045 | -21.8399 |  | 5.845 | -49.9666 | -22.8582 | -46.3522 | -26.0536
5.9725 | -8.0121 | -46.7118 | -21.6877 | -25.9131 | -11.4297 | -44.5671 | -25.9357 | -21.3887 |  | 5.9725 | -49.9064 | -22.5455 | -46.1167 | -25.9967
6.1 | -8.1616 | -46.3223 | -21.4384 | -25.7316 | -11.7704 | -45.3809 | -25.8628 | -20.9841 |  | 6.1 | -49.9081 | -22.2646 | -45.8199 | -25.9741
6.2275 | -8.3181 | -45.9787 | -21.2373 | -25.6213 | -12.1358 | -46.1731 | -25.7985 | -20.6757 |  | 6.2275 | -49.8789 | -22.0158 | -45.4672 | -26.0073
6.355 | -8.5 | -45.9175 | -21.0525 | -25.5375 | -12.5245 | -46.5397 | -25.7265 | -20.4637 |  | 6.355 | -49.8735 | -21.7811 | -45.0814 | -26.0456
6.4825 | -8.6014 | -46.0291 | -20.9644 | -25.4868 | -12.8228 | -46.7258 | -25.7071 | -20.2919 |  | 6.4825 | -49.8623 | -21.6084 | -44.7309 | -26.1255
6.61 | -8.7802 | -45.6556 | -20.802 | -25.4974 | -13.209 | -46.6538 | -25.6628 | -20.2007 |  | 6.61 | -49.9305 | -21.4733 | -44.2488 | -26.2027
6.7375 | -8.911 | -45.6092 | -20.7185 | -25.5361 | -13.5223 | -46.3095 | -25.6884 | -20.1168 |  | 6.7375 | -49.9085 | -21.3834 | -43.6416 | -26.2896
6.865 | -9.0036 | -45.5677 | -20.6536 | -25.5786 | -13.8942 | -45.8977 | -25.7476 | -20.0594 |  | 6.865 | -49.7707 | -21.3348 | -42.9753 | -26.3803
6.9925 | -9.0873 | -45.4946 | -20.62 | -25.6891 | -14.237 | -45.7157 | -25.8596 | -20.0309 |  | 6.9925 | -49.5804 | -21.3245 | -42.3233 | -26.48
7.12 | -9.1682 | -45.6798 | -20.637 | -25.848 | -14.4756 | -45.2761 | -26.0128 | -20.0281 |  | 7.12 | -49.3676 | -21.3495 | -41.7841 | -26.6101
7.2475 | -9.1714 | -46.0651 | -20.6935 | -26.0425 | -14.6919 | -44.9765 | -26.2036 | -20.0734 |  | 7.2475 | -49.1582 | -21.3964 | -41.3395 | -26.7593
7.375 | -9.1924 | -46.3515 | -20.7513 | -26.2369 | -14.9351 | -44.6927 | -26.4151 | -20.1432 |  | 7.375 | -48.9543 | -21.4819 | -41.02 | -26.9355
7.5025 | -9.1166 | -46.7999 | -20.8866 | -26.4322 | -14.978 | -44.1312 | -26.6754 | -20.152 |  | 7.5025 | -48.7133 | -21.5816 | -40.7294 | -27.1207
7.63 | -8.9706 | -47.5311 | -21.0148 | -26.6542 | -15.169 | -43.4459 | -26.9199 | -20.2693 |  | 7.63 | -48.4108 | -21.687 | -40.4812 | -27.3199
7.7575 | -8.8729 | -47.7972 | -21.1466 | -26.8881 | -15.3222 | -42.9812 | -27.2125 | -20.4191 |  | 7.7575 | -48.0478 | -21.8097 | -40.2592 | -27.5478
7.885 | -8.6961 | -48.059 | -21.322 | -27.124 | -15.5056 | -42.32 | -27.5014 | -20.5766 |  | 7.885 | -47.6379 | -21.9493 | -39.993 | -27.8
8.0125 | -8.5242 | -48.1433 | -21.5204 | -27.3752 | -15.6714 | -41.7664 | -27.7941 | -20.8199 |  | 8.0125 | -47.2066 | -22.0938 | -39.7621 | -28.0735
8.14 | -8.3546 | -48.0446 | -21.7454 | -27.6787 | -15.9244 | -41.501 | -28.0993 | -21.1816 |  | 8.14 | -46.7926 | -22.2629 | -39.6086 | -28.3857
8.2675 | -8.195 | -47.5328 | -22.0106 | -27.9919 | -16.1671 | -41.3377 | -28.4154 | -21.5442 |  | 8.2675 | -46.3637 | -22.4478 | -39.5175 | -28.7123
8.395 | -8.053 | -47.4263 | -22.3014 | -28.3282 | -16.4349 | -41.4263 | -28.6979 | -21.9821 |  | 8.395 | -45.8718 | -22.659 | -39.4468 | -29.0611
8.5225 | -7.9253 | -47.2941 | -22.6448 | -28.6192 | -16.627 | -41.7221 | -29.0255 | -22.4347 |  | 8.5225 | -45.4319 | -22.9 | -39.4169 | -29.4337
8.65 | -7.6965 | -47.219 | -23.0049 | -28.8921 | -16.9763 | -42.0723 | -29.299 | -22.8797 |  | 8.65 | -45.0398 | -23.1756 | -39.5099 | -29.8282
8.7775 | -7.657 | -47.5351 | -23.3875 | -29.1999 | -17.0567 | -42.5898 | -29.6401 | -23.3719 |  | 8.7775 | -44.7936 | -23.4824 | -39.641 | -30.2416
8.905 | -7.6237 | -47.7512 | -23.806 | -29.5051 | -16.9758 | -43.2612 | -30.0097 | -23.8675 |  | 8.905 | -44.6984 | -23.8222 | -39.7742 | -30.679
9.0325 | -7.6457 | -47.9027 | -24.1967 | -29.8248 | -16.8973 | -43.7843 | -30.36 | -24.3956 |  | 9.0325 | -44.6307 | -24.1891 | -40.0429 | -31.1512
9.16 | -7.7401 | -48.1911 | -24.6166 | -30.2094 | -16.6772 | -44.5156 | -30.7611 | -24.9709 |  | 9.16 | -44.583 | -24.5831 | -40.3254 | -31.6492
9.2875 | -7.9099 | -48.3334 | -25.0817 | -30.7062 | -16.2941 | -45.2939 | -31.241 | -25.5612 |  | 9.2875 | -44.4873 | -25.0091 | -40.6776 | -32.176
9.415 | -8.0835 | -48.2982 | -25.514 | -31.2098 | -15.9683 | -45.9587 | -31.7236 | -26.0994 |  | 9.415 | -44.4908 | -25.4594 | -41.125 | -32.7446
9.5425 | -8.2161 | -48.238 | -26.0096 | -31.7832 | -15.6387 | -46.4912 | -32.2426 | -26.6647 |  | 9.5425 | -44.5012 | -25.9089 | -41.565 | -33.3673
9.67 | -8.4058 | -48.1374 | -26.5004 | -32.4361 | -15.305 | -47.0877 | -32.8502 | -27.2023 |  | 9.67 | -44.5579 | -26.3732 | -42.0382 | -34.0328
9.7975 | -8.5992 | -48.1374 | -27.0017 | -33.1223 | -14.9771 | -47.5439 | -33.4397 | -27.7184 |  | 9.7975 | -44.6616 | -26.8518 | -42.5532 | -34.7419
9.925 | -8.7857 | -48.4479 | -27.5289 | -33.8037 | -14.6429 | -48.1126 | -34.1196 | -28.2365 |  | 9.925 | -44.7518 | -27.3465 | -43.0889 | -35.5009
10.0525 | -8.9288 | -48.9299 | -28.0878 | -34.5205 | -14.3103 | -48.9012 | -34.8398 | -28.8369 |  | 10.0525 | -44.7996 | -27.864 | -43.6367 | -36.3048
10.18 | -9.2126 | -49.4511 | -28.5828 | -35.2301 | -13.9319 | -49.826 | -35.6122 | -29.4925 |  | 10.18 | -44.8429 | -28.4095 | -44.2782 | -37.1266
10.3075 | -9.3131 | -49.9586 | -29.2064 | -35.9712 | -13.5188 | -50.783 | -36.4303 | -30.2269 |  | 10.3075 | -44.8682 | -28.9894 | -45.2655 | -37.9345
10.435 | -9.4163 | -50.3679 | -29.8719 | -36.7702 | -13.0179 | -51.5947 | -37.3459 | -30.9144 |  | 10.435 | -44.9503 | -29.6064 | -46.5038 | -38.6994
10.5625 | -9.5843 | -50.599 | -30.5238 | -37.7168 | -12.4086 | -52.2816 | -38.2772 | -31.6233 |  | 10.5625 | -45.0556 | -30.2657 | -47.2871 | -39.3841
10.69 | -9.7573 | -50.5697 | -31.2312 | -38.8181 | -11.831 | -52.9591 | -39.3232 | -32.2696 |  | 10.69 | -45.1559 | -30.9541 | -47.8562 | -39.9872
10.8175 | -9.936 | -50.5861 | -31.97 | -40.0525 | -11.13 | -53.4867 | -40.4598 | -32.7741 |  | 10.8175 | -45.2572 | -31.6562 | -48.2995 | -40.5009
10.945 | -10.2052 | -50.5006 | -32.6462 | -41.399 | -10.5047 | -53.7504 | -41.6799 | -33.2102 |  | 10.945 | -45.3464 | -32.3516 | -48.6099 | -40.9315
11.0725 | -10.6499 | -50.2591 | -33.2704 | -42.8683 | -9.9528 | -53.8048 | -43.011 | -33.6832 |  | 11.0725 | -45.4553 | -33 | -48.7216 | -41.293
11.2 | -11.2483 | -49.8249 | -33.8312 | -44.3476 | -9.391 | -53.7433 | -44.3838 | -34.0727 |  | 11.2 | -45.5356 | -33.5772 | -48.7258 | -41.585
11.3275 | -11.8812 | -49.3976 | -34.3507 | -45.8231 | -8.9 | -53.3375 | -45.6732 | -34.4185 |  | 11.3275 | -45.5802 | -34.0658 | -48.5767 | -41.8079
11.455 | -12.7261 | -48.9337 | -34.8457 | -47.1804 | -8.4715 | -52.8874 | -46.7519 | -34.7611 |  | 11.455 | -45.5693 | -34.48 | -48.3054 | -41.9659
11.5825 | -13.7629 | -48.6339 | -35.2977 | -48.0549 | -8.058 | -52.6596 | -47.3851 | -35.0647 |  | 11.5825 | -45.4952 | -34.8179 | -47.9388 | -42.0668
11.71 | -14.9392 | -48.4433 | -35.7302 | -48.2945 | -7.6713 | -52.4804 | -47.468 | -35.3678 |  | 11.71 | -45.3941 | -35.0824 | -47.4942 | -42.1057
11.8375 | -16.2009 | -48.3265 | -36.1429 | -47.9624 | -7.3108 | -52.2211 | -47.0717 | -35.6954 |  | 11.8375 | -45.247 | -35.2574 | -47.0206 | -42.0886
11.965 | -17.7078 | -48.2321 | -36.5041 | -47.0466 | -6.9953 | -52.1328 | -46.1888 | -35.9661 |  | 11.965 | -45.0837 | -35.3615 | -46.5905 | -42.0151
12.0925 | -19.5798 | -48.1309 | -36.8531 | -45.7568 | -6.7632 | -51.969 | -45.0485 | -36.2375 |  | 12.0925 | -44.946 | -35.3936 | -46.1191 | -41.8887
12.22 | -21.2515 | -48.0435 | -37.2093 | -44.515 | -6.4209 | -51.9535 | -43.9255 | -36.5529 |  | 12.22 | -44.9008 | -35.386 | -45.686 | -41.6957
12.3475 | -22.1609 | -48.0153 | -37.5365 | -43.3925 | -6.2259 | -52.1815 | -42.8225 | -36.8992 |  | 12.3475 | -44.9481 | -35.3322 | -45.2096 | -41.4459
12.475 | -22.4014 | -48.0762 | -37.8602 | -42.3803 | -6.0977 | -52.6456 | -41.8389 | -37.3232 |  | 12.475 | -44.9701 | -35.2305 | -44.6262 | -41.1279
12.6025 | -21.7929 | -48.2173 | -38.1287 | -41.491 | -5.8778 | -54.9692 | -41.0405 | -37.8084 |  | 12.6025 | -44.9977 | -35.0833 | -43.9688 | -40.7317
12.73 | -19.8167 | -48.6435 | -38.35 | -40.6856 | -5.5868 | -57.182 | -40.3545 | -38.2302 |  | 12.73 | -45.168 | -34.894 | -43.2397 | -40.2704
12.8575 | -17.7014 | -48.9626 | -38.2731 | -39.9363 | -5.5461 | -58.5831 | -39.697 | -38.5106 |  | 12.8575 | -45.4344 | -34.6612 | -42.4527 | -39.7779
12.985 | -15.7188 | -49.4764 | -38.0602 | -39.3083 | -5.3247 | -61.5268 | -39.1845 | -38.477 |  | 12.985 | -45.7496 | -34.3685 | -41.3529 | -39.2801
13.1125 | -13.9541 | -50.0402 | -37.5686 | -38.733 | -5.1547 | -63.4088 | -38.6822 | -38.1648 |  | 13.1125 | -45.9926 | -34.0169 | -40.0219 | -38.8146
13.24 | -12.5053 | -50.581 | -36.8577 | -38.2612 | -5.1768 | -62.4891 | -38.2035 | -37.634 |  | 13.24 | -46.1705 | -33.5841 | -39.188 | -38.4116
13.3675 | -11.6193 | -50.9978 | -35.9276 | -37.8615 | -5.2493 | -61.6462 | -37.785 | -37.0007 |  | 13.3675 | -46.3179 | -33.0853 | -38.6106 | -38.0807
13.495 | -10.6591 | -51.541 | -35.0378 | -37.4604 | -5.1837 | -61.0065 | -37.4305 | -36.2081 |  | 13.495 | -46.5088 | -32.5278 | -38.1759 | -37.8253
13.6225 | -9.9012 | -51.883 | -34.0741 | -37.1363 | -5.3237 | -58.1574 | -37.1225 | -35.4303 |  | 13.6225 | -46.7182 | -31.9233 | -37.9087 | -37.6611
13.75 | -9.2029 | -52.1978 | -33.159 | -36.863 | -5.4289 | -56.2644 | -36.8929 | -34.6134 |  | 13.75 | -46.9119 | -31.3108 | -37.9095 | -37.5846
13.8775 | -8.6224 | -52.4784 | -32.302 | -36.5914 | -5.4145 | -55.3566 | -36.719 | -33.7407 |  | 13.8775 | -47.1264 | -30.7095 | -38.1163 | -37.5842
14.005 | -8.1283 | -52.6406 | -31.5041 | -36.4137 | -5.556 | -54.1861 | -36.6074 | -32.8398 |  | 14.005 | -47.4067 | -30.1475 | -38.5766 | -37.6868
14.1325 | -7.6898 | -52.7855 | -30.772 | -36.2067 | -5.6208 | -53.9136 | -36.543 | -31.972 |  | 14.1325 | -47.66 | -29.614 | -39.215 | -37.8912
14.26 | -7.2913 | -52.9523 | -30.0941 | -36.1236 | -5.5788 | -54.3714 | -36.5121 | -31.1109 |  | 14.26 | -47.9963 | -29.1138 | -39.9295 | -38.1988
14.3875 | -6.9106 | -53.1097 | -29.5263 | -36.1249 | -5.5717 | -54.6078 | -36.5377 | -30.3171 |  | 14.3875 | -48.4909 | -28.6563 | -40.6888 | -38.6172
14.515 | -6.4685 | -53.3456 | -29.0535 | -36.2155 | -5.5553 | -55.3274 | -36.5919 | -29.624 |  | 14.515 | -49.1341 | -28.26 | -41.5109 | -39.1001
14.6425 | -6.0611 | -53.6907 | -28.6143 | -36.302 | -5.3626 | -57.8297 | -36.6228 | -28.9939 |  | 14.6425 | -49.7503 | -27.8835 | -42.251 | -39.7052
14.77 | -5.6678 | -54.3507 | -28.2329 | -36.4963 | -5.3582 | -59.512 | -36.7314 | -28.4502 |  | 14.77 | -50.3056 | -27.5607 | -42.9471 | -40.3295
14.8975 | -5.3211 | -55.0951 | -27.8737 | -36.6216 | -5.4148 | -60.7315 | -36.8721 | -27.9894 |  | 14.8975 | -50.7781 | -27.266 | -43.6508 | -40.9714
15.025 | -5.0134 | -56.2476 | -27.5272 | -36.7759 | -5.4657 | -62.1998 | -37.0134 | -27.5828 |  | 15.025 | -51.1381 | -27.0063 | -44.3653 | -41.912
15.1525 | -4.8153 | -57.584 | -27.2299 | -36.9381 | -5.6129 | -62.6136 | -37.2131 | -27.2574 |  | 15.1525 | -51.6086 | -26.7744 | -45.2202 | -42.8597
15.28 | -4.61 | -59.4956 | -26.9517 | -37.156 | -5.8573 | -61.3945 | -37.438 | -26.9503 |  | 15.28 | -52.0982 | -26.5582 | -46.1584 | -43.788
15.4075 | -4.4724 | -61.9467 | -26.6969 | -37.3835 | -5.8554 | -60.0819 | -37.4956 | -26.6953 |  | 15.4075 | -52.4539 | -26.3699 | -47.1157 | -44.3965
15.535 | -4.3412 | -64.4824 | -26.4898 | -37.6844 | -5.9627 | -58.7319 | -37.6671 | -26.5198 |  | 15.535 | -52.7023 | -26.1785 | -48.001 | -44.7947
15.6625 | -4.28 | -66.9333 | -26.3137 | -38.0308 | -6.135 | -57.1838 | -37.8513 | -26.3349 |  | 15.6625 | -53.0528 | -25.99 | -48.8831 | -45.1073
15.79 | -4.0851 | -70.2622 | -26.0492 | -38.3012 | -6.0902 | -55.7687 | -37.9051 | -26.1935 |  | 15.79 | -53.7246 | -25.8414 | -49.6517 | -45.3186
15.9175 | -4.0024 | -71.6951 | -25.9154 | -38.5724 | -6.1357 | -54.5288 | -38.0664 | -26.0815 |  | 15.9175 | -54.6842 | -25.7214 | -50.4169 | -45.4881
16.045 | -3.9379 | -72.2509 | -25.8087 | -39.0435 | -6.2855 | -53.6106 | -38.2916 | -25.9945 |  | 16.045 | -55.945 | -25.6412 | -51.1213 | -45.5987
16.1725 | -3.8893 | -71.664 | -25.7389 | -39.4889 | -6.2524 | -52.9884 | -38.446 | -25.9127 |  | 16.1725 | -56.7126 | -25.5928 | -51.7411 | -45.6637
16.3 | -3.8255 | -70.4597 | -25.7345 | -39.8176 | -6.1009 | -52.4656 | -38.507 | -25.9062 |  | 16.3 | -57.1602 | -25.5729 | -52.2521 | -45.6411
16.4275 | -3.9065 | -67.6434 | -25.8617 | -40.6526 | -6.4529 | -52.1195 | -39.1098 | -25.8901 |  | 16.4275 | -57.4458 | -25.5911 | -52.5983 | -45.5215
16.555 | -3.9136 | -65.9613 | -25.8612 | -41.94 | -6.7161 | -51.7482 | -39.6826 | -25.868 |  | 16.555 | -57.662 | -25.6372 | -52.7581 | -45.3266
16.6825 | -3.922 | -65.0244 | -26.0549 | -43.3846 | -7.2115 | -51.4273 | -40.6077 | -25.8794 |  | 16.6825 | -57.8001 | -25.7049 | -52.7585 | -45.0404
16.81 | -3.96 | -64.0702 | -26.2365 | -45.2354 | -7.6602 | -51.0626 | -41.515 | -25.9195 |  | 16.81 | -57.9537 | -25.7862 | -52.6676 | -44.6774
16.9375 | -3.9966 | -62.8443 | -26.4728 | -47.475 | -8.1878 | -50.7114 | -42.5953 | -25.9584 |  | 16.9375 | -58.0386 | -25.8985 | -52.5982 | -44.2014
17.065 | -4.0626 | -62.1938 | -26.565 | -49.5319 | -8.6167 | -50.4763 | -43.7648 | -25.963 |  | 17.065 | -58.0172 | -26.0261 | -52.6082 | -43.6248
17.1925 | -4.2091 | -61.6265 | -26.6256 | -51.1636 | -9.4016 | -50.3537 | -45.8137 | -25.9842 |  | 17.1925 | -57.8923 | -26.1814 | -52.8616 | -42.9922
17.32 | -4.3843 | -60.2579 | -26.4792 | -52.0382 | -10.2861 | -50.233 | -48.1321 | -25.9737 |  | 17.32 | -57.7925 | -26.3652 | -53.2068 | -42.2745
17.4475 | -4.5503 | -59.5995 | -26.3055 | -51.752 | -11.1522 | -50.1064 | -50.3082 | -25.9062 |  | 17.4475 | -57.7523 | -26.5393 | -53.6105 | -41.5497
17.575 | -4.7917 | -59.394 | -26.0632 | -50.7996 | -12.2835 | -49.9931 | -51.709 | -25.8659 |  | 17.575 | -57.7298 | -26.7264 | -54.142 | -40.8129
17.7025 | -4.9606 | -58.8541 | -25.9456 | -49.1849 | -12.8435 | -49.6237 | -53.1471 | -25.8299 |  | 17.7025 | -57.6376 | -26.9129 | -54.5793 | -39.7726
17.83 | -5.0822 | -58.0685 | -25.9033 | -47.9497 | -13.2619 | -49.2662 | -52.7144 | -25.8506 |  | 17.83 | -57.3725 | -27.0982 | -54.6721 | -38.7171
17.9575 | -5.3434 | -57.6933 | -25.83 | -46.9184 | -14.1244 | -48.9223 | -50.8364 | -25.8288 |  | 17.9575 | -57.0708 | -27.29 | -54.487 | -37.6768
18.085 | -5.6023 | -57.4105 | -25.7575 | -45.9923 | -14.9058 | -48.5562 | -48.9164 | -25.7852 |  | 18.085 | -56.8002 | -27.4723 | -54.1813 | -36.966
18.2125 | -6.0043 | -57.3345 | -25.6207 | -45.0074 | -15.5253 | -48.2196 | -47.3131 | -25.6908 |  | 18.2125 | -56.5319 | -27.667 | -53.7782 | -36.4471
18.34 | -6.5178 | -57.4021 | -25.5134 | -44.1911 | -16.4186 | -48.0971 | -44.9545 | -25.6471 |  | 18.34 | -56.1233 | -27.89 | -53.3236 | -36.0155
18.4675 | -7.055 | -57.7987 | -25.3917 | -42.725 | -16.2646 | -47.8284 | -43.6198 | -25.559 |  | 18.4675 | -55.4497 | -28.0954 | -52.8951 | -35.662
18.595 | -7.4041 | -57.7228 | -25.2927 | -41.3666 | -15.2793 | -47.5238 | -42.8999 | -25.5476 |  | 18.595 | -54.4846 | -28.3102 | -52.4511 | -35.3331
18.7225 | -8.0067 | -57.9073 | -25.2938 | -40.4991 | -14.4114 | -47.3157 | -42.0789 | -25.621 |  | 18.7225 | -53.2845 | -28.5186 | -52.009 | -35.0429
18.85 | -8.6347 | -57.8295 | -25.3339 | -39.7805 | -13.38 | -47.1764 | -41.2821 | -25.6706 |  | 18.85 | -52.6104 | -28.7439 | -51.6054 | -34.7847
18.9775 | -9.4064 | -57.5494 | -25.4089 | -39.1509 | -12.1896 | -47.0061 | -40.5123 | -25.7829 |  | 18.9775 | -52.3374 | -28.9986 | -51.2723 | -34.5882
19.105 | -10.2369 | -57.2629 | -25.5245 | -38.5702 | -11.6237 | -46.9902 | -39.47 | -25.9258 |  | 19.105 | -52.4212 | -29.2698 | -51.0182 | -34.4524
19.2325 | -11.2452 | -57.6491 | -25.6118 | -37.9551 | -10.9255 | -47.0015 | -38.5261 | -25.9944 |  | 19.2325 | -52.6891 | -29.563 | -50.8558 | -34.3677
19.36 | -12.2932 | -57.5795 | -25.7061 | -37.316 | -10.1074 | -46.9934 | -37.7034 | -26.1006 |  | 19.36 | -52.825 | -29.8723 | -50.716 | -34.2969
19.4875 | -13.4141 | -57.7659 | -25.8266 | -36.7694 | -9.1973 | -46.959 | -37.0861 | -26.2618 |  | 19.4875 | -52.8958 | -30.2033 | -50.6038 | -34.2073
19.615 | -14.737 | -58.2602 | -25.9043 | -36.2157 | -8.3925 | -47.0488 | -36.4608 | -26.3366 |  | 19.615 | -52.9427 | -30.5754 | -50.4488 | -34.1344
19.7425 | -16.8469 | -58.1714 | -26.0875 | -35.7414 | -7.7037 | -47.2381 | -35.8843 | -26.4224 |  | 19.7425 | -52.8364 | -31.1283 | -50.1507 | -34.0378
19.87 | -19.6765 | -57.222 | -26.3714 | -35.2454 | -7.1733 | -47.5745 | -35.3568 | -26.7164 |  | 19.87 | -52.6156 | -31.8588 | -49.5729 | -33.9065
19.9975 | -23.2434 | -56.7389 | -26.5859 | -34.7237 | -6.7269 | -48.038 | -34.8295 | -27.0414 |  | 19.9975 | -52.3395 | -32.2191 | -48.9324 | -33.8112
20.125 | -25.1067 | -56.232 | -26.8291 | -34.2545 | -6.2947 | -48.715 | -34.3847 | -27.4682 |  | 20.125 | -51.973 | -32.552 | -48.233 | -34.1223
20.2525 | -26.0966 | -55.3415 | -27.1833 | -33.7982 | -5.9348 | -49.3668 | -33.9595 | -27.9655 |  | 20.2525 | -51.652 | -32.8621 | -47.4163 | -34.4566
20.38 | -25.4731 | -55.1147 | -27.4566 | -33.3075 | -5.585 | -49.9071 | -33.5325 | -28.4875 |  | 20.38 | -51.386 | -33.1692 | -46.7203 | -34.6606
20.5075 | -23.4751 | -55.1046 | -27.8411 | -32.8489 | -5.303 | -50.4086 | -33.1176 | -29.0941 |  | 20.5075 | -51.0997 | -33.4876 | -46.2185 | -34.8023
20.635 | -20.1535 | -55.0159 | -28.333 | -32.3596 | -5.0639 | -50.6455 | -32.6678 | -29.7612 |  | 20.635 | -50.8137 | -33.8214 | -45.8731 | -34.9119
20.7625 | -17.9905 | -54.6819 | -29.0971 | -31.7579 | -4.9182 | -50.4674 | -32.0831 | -30.7375 |  | 20.7625 | -50.5174 | -34.1582 | -45.596 | -34.989
20.89 | -15.9925 | -53.9779 | -31.0463 | -31.1633 | -4.81 | -50.2939 | -31.5575 | -32.9292 |  | 20.89 | -50.2284 | -34.4973 | -45.4366 | -35.0563
21.0175 | -14.4971 | -52.8322 | -32.1132 | -30.9255 | -4.7548 | -50.6148 | -31.5469 | -33.9086 |  | 21.0175 | -49.9289 | -34.8312 | -45.293 | -35.1001
21.145 | -13.3089 | -51.6364 | -32.8662 | -32.6039 | -4.6405 | -51.0707 | -33.9519 | -34.4561 |  | 21.145 | -49.6167 | -35.1655 | -45.142 | -35.1325
21.2725 | -12.1898 | -50.4432 | -33.4691 | -34.3332 | -4.5105 | -51.7188 | -35.6481 | -34.7971 |  | 21.2725 | -49.3392 | -35.5026 | -44.9259 | -35.1492
21.4 | -11.1713 | -49.4213 | -33.7617 | -35.7249 | -4.3591 | -52.5311 | -36.9037 | -34.695 |  | 21.4 | -49.0169 | -35.8344 | -44.7004 | -35.1421
21.5275 | -10.2523 | -49.1861 | -32.8332 | -36.9644 | -4.1954 | -53.3336 | -37.9888 | -33.344 |  | 21.5275 | -48.6075 | -36.1435 | -44.4485 | -35.109
21.655 | -9.4832 | -49.2846 | -32.8455 | -37.747 | -4.0186 | -53.6599 | -38.5132 | -33.1527 |  | 21.655 | -48.0923 | -36.4047 | -44.1599 | -35.0416
21.7825 | -8.8096 | -49.529 | -33.1246 | -36.5969 | -3.9255 | -53.6702 | -36.6042 | -33.1277 |  | 21.7825 | -47.36 | -36.6142 | -43.8052 | -34.9207
21.91 | -8.3384 | -49.7654 | -33.4516 | -35.4979 | -3.8724 | -53.6094 | -35.4809 | -33.17 |  | 21.91 | -46.4932 | -36.7624 | -43.4169 | -34.748
22.0375 | -7.9572 | -50.0289 | -33.7692 | -34.8588 | -3.8127 | -53.5307 | -34.9002 | -33.3123 |  | 22.0375 | -45.7632 | -36.8458 | -43.0159 | -34.5344
22.165 | -7.706 | -50.2692 | -34.1331 | -34.4071 | -3.7829 | -53.384 | -34.5061 | -33.3966 |  | 22.165 | -45.0845 | -36.8644 | -42.575 | -34.288
22.2925 | -7.576 | -50.4687 | -34.2668 | -34.1348 | -3.7257 | -53.1586 | -34.212 | -33.5857 |  | 22.2925 | -44.4076 | -36.7666 | -42.1331 | -34.007
22.42 | -7.5222 | -50.6387 | -34.2561 | -33.9396 | -3.6482 | -52.9886 | -33.987 | -33.7673 |  | 22.42 | -43.8383 | -36.4191 | -41.7613 | -33.7034
22.5475 | -7.495 | -50.797 | -34.2316 | -33.7548 | -3.5955 | -52.7144 | -33.785 | -33.8246 |  | 22.5475 | -43.385 | -35.8213 | -41.3718 | -33.38
22.675 | -7.4999 | -50.9091 | -34.2292 | -33.6547 | -3.5778 | -52.3096 | -33.6454 | -33.8225 |  | 22.675 | -43.0384 | -35.5254 | -41.0152 | -32.9514
22.8025 | -7.6273 | -50.9483 | -33.9521 | -33.5366 | -3.5438 | -51.8552 | -33.4787 | -33.7533 |  | 22.8025 | -42.8051 | -35.2114 | -40.6886 | -32.0708
22.93 | -7.6338 | -51.0583 | -33.8483 | -33.6952 | -3.5341 | -51.3345 | -33.3049 | -33.63 |  | 22.93 | -42.5888 | -34.8673 | -40.3146 | -31.1075
23.0575 | -7.7494 | -51.0744 | -33.6107 | -33.8374 | -3.5554 | -50.7225 | -33.0923 | -33.3872 |  | 23.0575 | -42.4348 | -34.4787 | -39.9233 | -30.24
23.185 | -7.8908 | -51.1053 | -33.2242 | -33.9574 | -3.5705 | -50.1217 | -32.7814 | -33.154 |  | 23.185 | -42.3713 | -34.0403 | -39.5678 | -29.4098
23.3125 | -8.075 | -51.0803 | -32.709 | -33.9088 | -3.5838 | -49.4867 | -32.3173 | -32.7598 |  | 23.3125 | -42.3133 | -33.5449 | -39.2812 | -28.5844
23.44 | -8.1743 | -50.985 | -32.3573 | -33.6609 | -3.6066 | -48.7824 | -31.7474 | -32.3624 |  | 23.44 | -42.2441 | -33.0058 | -39.0551 | -27.7624
23.5675 | -8.5519 | -50.8096 | -31.5747 | -32.869 | -3.6245 | -48.0738 | -30.9777 | -31.6472 |  | 23.5675 | -42.1492 | -32.4298 | -38.8701 | -26.948
23.695 | -8.8436 | -50.4692 | -30.9298 | -31.7967 | -3.6929 | -47.3268 | -30.0186 | -30.9108 |  | 23.695 | -42.0473 | -31.8158 | -38.7322 | -26.1419
23.8225 | -9.3821 | -50.0109 | -30.1438 | -30.5106 | -3.7611 | -46.5385 | -28.91 | -30.0282 |  | 23.8225 | -41.9464 | -31.1708 | -38.6668 | -25.3538
23.95 | -9.9996 | -49.4399 | -29.2383 | -28.8869 | -3.8079 | -45.756 | -27.6151 | -29.0747 |  | 23.95 | -41.8438 | -30.4939 | -38.6777 | -24.5903
24.0775 | -10.7273 | -48.6952 | -28.2249 | -27.3051 | -3.8849 | -44.9747 | -26.1996 | -28.0713 |  | 24.0775 | -41.7212 | -29.7994 | -38.7448 | -23.8555
24.205 | -11.4559 | -47.6879 | -27.4707 | -25.6481 | -3.9907 | -44.2026 | -24.7505 | -27.1425 |  | 24.205 | -41.5953 | -29.1011 | -38.8216 | -23.1512
24.3325 | -12.3539 | -46.7559 | -26.5347 | -24.0217 | -4.0751 | -43.4981 | -23.2677 | -26.1775 |  | 24.3325 | -41.4541 | -28.4159 | -38.9084 | -22.4704
24.46 | -13.3183 | -45.7414 | -25.7391 | -22.3811 | -4.1725 | -42.9107 | -21.8343 | -25.2986 |  | 24.46 | -41.3369 | -27.7443 | -39.0078 | -21.8314
24.5875 | -14.353 | -44.6848 | -24.9785 | -20.9432 | -4.3083 | -42.3967 | -20.4853 | -24.4793 |  | 24.5875 | -41.2262 | -27.0995 | -39.0703 | -21.2264
24.715 | -15.4831 | -43.7861 | -24.1189 | -19.556 | -4.432 | -41.9854 | -19.2136 | -23.7364 |  | 24.715 | -41.1118 | -26.4766 | -39.1257 | -20.6632
24.8425 | -16.4825 | -43.1661 | -23.2746 | -18.3788 | -4.5651 | -41.7587 | -18.0998 | -23.0763 |  | 24.8425 | -40.992 | -25.8816 | -39.1844 | -20.15
24.97 | -17.0456 | -42.5546 | -22.4723 | -17.3523 | -4.6894 | -41.506 | -17.114 | -22.4506 |  | 24.97 | -40.8431 | -25.3231 | -39.2057 | -19.6873
25.0975 | -16.9326 | -42.1106 | -21.6972 | -16.5085 | -4.8235 | -41.2641 | -16.2451 | -21.8318 |  | 25.0975 | -40.6347 | -24.802 | -39.1668 | -19.2689
25.225 | -16.3089 | -41.8046 | -21.011 | -15.8558 | -4.9638 | -41.1421 | -15.5543 | -21.2691 |  | 25.225 | -40.363 | -24.3048 | -39.1122 | -18.8972
25.3525 | -15.2888 | -41.4098 | -20.3973 | -15.3006 | -5.0902 | -40.886 | -14.9984 | -20.6576 |  | 25.3525 | -40.0516 | -23.8548 | -39.0337 | -18.5672
25.48 | -13.9125 | -40.9371 | -19.8322 | -14.8382 | -5.2379 | -40.4573 | -14.5113 | -20.0836 |  | 25.48 | -39.7002 | -23.4371 | -38.9494 | -18.2727
25.6075 | -12.5807 | -40.4906 | -19.3977 | -14.5271 | -5.3917 | -40.0902 | -14.2071 | -19.6412 |  | 25.6075 | -39.3343 | -23.0559 | -38.8722 | -18.0345
25.735 | -11.395 | -39.7401 | -19.0236 | -14.2805 | -5.5409 | -39.465 | -14.0109 | -19.2594 |  | 25.735 | -38.9516 | -22.7046 | -38.7812 | -17.8343
25.8625 | -10.3737 | -38.879 | -18.7381 | -14.145 | -5.7009 | -38.658 | -13.9269 | -18.9641 |  | 25.8625 | -38.5634 | -22.3833 | -38.6894 | -17.6729
25.99 | -9.4374 | -37.9756 | -18.5171 | -14.2047 | -5.84 | -37.8286 | -13.9256 | -18.781 |  | 25.99 | -38.1937 | -22.0972 | -38.5923 | -17.5548
26.1175 | -8.6914 | -36.9202 | -18.3426 | -14.2644 | -5.954 | -36.8924 | -13.9824 | -18.6303 |  | 26.1175 | -37.8407 | -21.8413 | -38.4766 | -17.4624
26.245 | -8.075 | -35.8783 | -18.2341 | -14.3403 | -6.0663 | -35.9159 | -14.0644 | -18.5371 |  | 26.245 | -37.5195 | -21.6189 | -38.3347 | -17.3908
26.3725 | -7.6404 | -35.1387 | -18.1869 | -14.4563 | -6.1408 | -35.2194 | -14.1525 | -18.5208 |  | 26.3725 | -37.2305 | -21.4293 | -38.1778 | -17.3383
26.5 | -7.365 | -34.5717 | -18.1431 | -14.5329 | -6.1838 | -34.6946 | -14.1984 | -18.5053 |  | 26.5 | -36.9603 | -21.2642 | -38.0081 | -17.2963

## IFCL IFRL
IF (GHz) | High Side LO | 8 GHz RF "A" | (8 GHz RF) IF RL | High Side LO | 8 GHz RF "B" | (8 GHz RF) IF RL | Low Side LO | 18 GHz RF "A" | (18 GHz RF) IF RL | Low Side LO | 18 GHz RF "B" | (18 GHz RF) IF RL
0.01 | -7.5187 | -0.13738 | -21.078 | -8.2648 | 0 | -20.1951 | -7.607 | -0.35467 | -13.3496 | -8.9807 | -0.21772 | -26.6797
0.0799 | -7.4523 | -0.071018 | -20.9799 | -8.3349 | -0.07015 | -19.7971 | -7.556 | -0.30372 | -13.5536 | -8.9282 | -0.16524 | -26.7934
0.1498 | -7.3813 | 0 | -20.7942 | -8.3573 | -0.092493 | -19.5862 | -7.4998 | -0.24755 | -13.7026 | -8.8763 | -0.1134 | -26.5753
0.2197 | -7.4232 | -0.041864 | -20.67 | -8.3383 | -0.073545 | -19.2522 | -7.4888 | -0.23654 | -13.8881 | -8.851 | -0.088041 | -26.3466
0.2896 | -7.4318 | -0.050537 | -20.4992 | -8.3037 | -0.03892 | -19.0002 | -7.5031 | -0.25085 | -13.9376 | -8.8216 | -0.058608 | -25.5322
0.3595 | -7.4599 | -0.078651 | -20.3309 | -8.3812 | -0.11638 | -19.3688 | -7.535 | -0.28266 | -14.0075 | -8.8161 | -0.053183 | -24.278
0.4294 | -7.4845 | -0.10324 | -20.2652 | -8.3958 | -0.131 | -19.4799 | -7.4924 | -0.24015 | -14.1193 | -8.7852 | -0.022225 | -22.956
0.4993 | -7.5069 | -0.12558 | -20.1231 | -8.4165 | -0.15166 | -19.4076 | -7.479 | -0.22671 | -14.468 | -8.784 | -0.021049 | -21.5992
0.5692 | -7.5706 | -0.18928 | -19.7311 | -8.3963 | -0.13154 | -19.4726 | -7.4172 | -0.16491 | -14.8121 | -8.7675 | -0.0045367 | -20.1974
0.6391 | -7.6185 | -0.23721 | -19.649 | -8.4609 | -0.19608 | -19.7264 | -7.4041 | -0.15177 | -15.2049 | -8.7629 | 0 | -19.1413
0.709 | -7.6714 | -0.29012 | -19.5709 | -8.5434 | -0.27863 | -19.8274 | -7.3698 | -0.11747 | -15.2456 | -8.7744 | -0.011467 | -18.395
0.7789 | -7.6803 | -0.299 | -19.3269 | -8.5952 | -0.33039 | -19.6194 | -7.3724 | -0.12011 | -15.442 | -8.785 | -0.022069 | -17.5051
0.8488 | -7.7182 | -0.33688 | -19.0117 | -8.6789 | -0.41409 | -19.5909 | -7.3342 | -0.081884 | -15.2003 | -8.8125 | -0.0496 | -16.9633
0.9187 | -7.7976 | -0.41632 | -19.0944 | -8.7226 | -0.45777 | -19.2582 | -7.3158 | -0.063522 | -15.1075 | -8.8253 | -0.062332 | -16.2814
0.9886 | -7.8605 | -0.47922 | -18.626 | -8.822 | -0.55721 | -19.165 | -7.2701 | -0.017806 | -14.9697 | -8.8325 | -0.069552 | -15.7888
1.0585 | -7.9448 | -0.56352 | -18.3817 | -8.8843 | -0.61954 | -18.7481 | -7.2653 | -0.012972 | -15.0506 | -8.8203 | -0.057399 | -15.2215
1.1284 | -7.9891 | -0.60778 | -18.4281 | -8.9943 | -0.72946 | -18.4131 | -7.2523 | 0 | -14.8601 | -8.8176 | -0.054674 | -14.9016
1.1983 | -8.0057 | -0.62445 | -18.5787 | -9.0519 | -0.78708 | -17.7371 | -7.2859 | -0.033553 | -14.9069 | -8.8352 | -0.072293 | -14.5557
1.2682 | -7.9863 | -0.60501 | -18.5649 | -9.1403 | -0.8755 | -17.2709 | -7.3056 | -0.053313 | -14.6878 | -8.8761 | -0.11312 | -14.3255
1.3381 | -7.9753 | -0.59397 | -18.9726 | -9.1998 | -0.93502 | -16.668 | -7.3255 | -0.073155 | -14.5829 | -8.9311 | -0.1682 | -14.0192
1.408 | -7.9717 | -0.59042 | -19.1424 | -9.2542 | -0.98938 | -16.1248 | -7.3281 | -0.075829 | -14.4304 | -8.954 | -0.19109 | -13.7515
1.4779 | -7.9933 | -0.61205 | -19.0971 | -9.2868 | -1.022 | -15.6108 | -7.3176 | -0.065295 | -14.2742 | -8.9937 | -0.23076 | -13.4539
1.5478 | -7.9992 | -0.61793 | -19.325 | -9.3082 | -1.0434 | -15.2048 | -7.3066 | -0.054307 | -14.0033 | -8.9889 | -0.22591 | -13.1504
1.6177 | -8.0254 | -0.64412 | -19.5983 | -9.3723 | -1.1075 | -14.7735 | -7.3115 | -0.059213 | -13.872 | -9.0315 | -0.26854 | -12.9198
1.6876 | -8.0227 | -0.64137 | -19.579 | -9.4117 | -1.1469 | -14.4281 | -7.3361 | -0.08378 | -13.5773 | -9.0529 | -0.28996 | -12.7035
1.7575 | -8.0264 | -0.64511 | -19.5846 | -9.4334 | -1.1686 | -14.0661 | -7.4048 | -0.15253 | -13.3298 | -9.1216 | -0.35866 | -12.5164
1.8274 | -8.0476 | -0.66636 | -19.6156 | -9.4154 | -1.1507 | -13.7058 | -7.4071 | -0.15479 | -13.2236 | -9.1346 | -0.37164 | -12.3658
1.8973 | -8.0739 | -0.6926 | -19.3143 | -9.4769 | -1.2121 | -13.4023 | -7.4241 | -0.17175 | -13.1465 | -9.1831 | -0.42018 | -12.2513
1.9672 | -8.1032 | -0.72192 | -18.924 | -9.5149 | -1.2501 | -13.2435 | -7.404 | -0.15166 | -13.0248 | -9.2085 | -0.44557 | -12.0499
2.0371 | -8.1262 | -0.74487 | -18.6145 | -9.5654 | -1.3007 | -12.9832 | -7.4207 | -0.16842 | -12.9632 | -9.2625 | -0.49959 | -11.8555
2.107 | -8.1515 | -0.77022 | -18.1102 | -9.5468 | -1.282 | -12.7912 | -7.4848 | -0.23245 | -12.8391 | -9.3336 | -0.57068 | -11.6847
2.1769 | -8.2154 | -0.83409 | -17.6062 | -9.5886 | -1.3239 | -12.6502 | -7.4893 | -0.23698 | -12.6072 | -9.3884 | -0.6255 | -11.5012
2.2468 | -8.2644 | -0.88307 | -17.1595 | -9.6853 | -1.4205 | -12.5236 | -7.5491 | -0.29676 | -12.4421 | -9.4733 | -0.71039 | -11.324
2.3167 | -8.2804 | -0.89914 | -16.5133 | -9.7511 | -1.4864 | -12.3865 | -7.5818 | -0.32954 | -12.2507 | -9.5472 | -0.78423 | -11.1628
2.3866 | -8.2819 | -0.90065 | -15.8818 | -9.7685 | -1.5037 | -12.312 | -7.6425 | -0.39022 | -12.0851 | -9.63 | -0.86707 | -11.0339
2.4565 | -8.3084 | -0.92708 | -15.3841 | -9.7552 | -1.4904 | -12.265 | -7.698 | -0.44572 | -11.8596 | -9.7061 | -0.94318 | -10.909
2.5264 | -8.3521 | -0.97085 | -14.739 | -9.7606 | -1.4958 | -12.2138 | -7.7434 | -0.49107 | -11.6067 | -9.7708 | -1.0079 | -10.7634
2.5963 | -8.3806 | -0.9993 | -14.1226 | -9.8004 | -1.5356 | -12.0761 | -7.8299 | -0.5776 | -11.3158 | -9.835 | -1.0721 | -10.6523
2.6662 | -8.3978 | -1.0165 | -13.6785 | -9.7839 | -1.5192 | -11.8821 | -7.9261 | -0.67382 | -11.0255 | -9.8957 | -1.1328 | -10.5821
2.7361 | -8.3919 | -1.0106 | -13.1877 | -9.7963 | -1.5315 | -11.6904 | -8.0023 | -0.75001 | -10.676 | -9.9373 | -1.1743 | -10.4703
2.806 | -8.3963 | -1.015 | -12.8014 | -9.7875 | -1.5227 | -11.4206 | -8.1118 | -0.85954 | -10.3548 | -9.978 | -1.215 | -10.3392
2.8759 | -8.3781 | -0.99679 | -12.5615 | -9.8412 | -1.5764 | -11.1426 | -8.2206 | -0.96825 | -10.1266 | -10.038 | -1.2751 | -10.1546
2.9458 | -8.3514 | -0.97013 | -12.3385 | -9.8767 | -1.6119 | -10.9164 | -8.3139 | -1.0616 | -9.8785 | -10.0815 | -1.3186 | -9.9676
3.0157 | -8.3262 | -0.94494 | -12.1098 | -9.8958 | -1.631 | -10.6916 | -8.4002 | -1.1479 | -9.6223 | -10.1506 | -1.3876 | -9.799
3.0856 | -8.3063 | -0.925 | -11.9119 | -9.9043 | -1.6395 | -10.434 | -8.4504 | -1.1981 | -9.4095 | -10.1853 | -1.4224 | -9.6023
3.1555 | -8.3619 | -0.98059 | -11.6387 | -9.9372 | -1.6724 | -10.1989 | -8.5559 | -1.3036 | -9.1429 | -10.236 | -1.4731 | -9.4286
3.2254 | -8.4226 | -1.0413 | -11.2949 | -10.0038 | -1.739 | -10.0082 | -8.6411 | -1.3888 | -8.8889 | -10.2987 | -1.5358 | -9.2926
3.2953 | -8.4795 | -1.0982 | -10.9388 | -10.0561 | -1.7913 | -9.7541 | -8.7234 | -1.4711 | -8.6974 | -10.3827 | -1.6198 | -9.1029
3.3652 | -8.5456 | -1.1643 | -10.519 | -10.1157 | -1.8509 | -9.5095 | -8.7498 | -1.4975 | -8.5531 | -10.4222 | -1.6592 | -8.8746
3.4351 | -8.5972 | -1.2159 | -10.1553 | -10.179 | -1.9142 | -9.2378 | -8.727 | -1.4747 | -8.4359 | -10.4836 | -1.7207 | -8.6593
3.505 | -8.7172 | -1.3359 | -9.8717 | -10.2782 | -2.0134 | -8.9438 | -8.7398 | -1.4875 | -8.3459 | -10.4748 | -1.7119 | -8.4426
3.5749 | -8.7979 | -1.4166 | -9.5177 | -10.3809 | -2.1161 | -8.629 | -8.8084 | -1.5561 | -8.2026 | -10.5645 | -1.8016 | -8.2706
3.6448 | -8.8919 | -1.5107 | -9.2355 | -10.4772 | -2.2124 | -8.3269 | -8.9645 | -1.7122 | -8.0816 | -10.6448 | -1.8819 | -8.0856
3.7147 | -8.9745 | -1.5932 | -8.9377 | -10.5816 | -2.3168 | -7.9688 | -9.0268 | -1.7745 | -7.9126 | -10.7489 | -1.9859 | -7.9303
3.7846 | -9.0971 | -1.7158 | -8.655 | -10.6986 | -2.4338 | -7.6685 | -9.0891 | -1.8368 | -7.7725 | -10.8764 | -2.1135 | -7.7696
3.8545 | -9.2125 | -1.8312 | -8.3004 | -10.8453 | -2.5805 | -7.3639 | -9.115 | -1.8627 | -7.6186 | -10.9549 | -2.192 | -7.6099
3.9244 | -9.3435 | -1.9622 | -8.0432 | -10.9847 | -2.7199 | -7.0773 | -9.2519 | -1.9996 | -7.4886 | -11.1353 | -2.3723 | -7.3795
3.9943 | -9.4316 | -2.0503 | -7.7026 | -11.122 | -2.8572 | -6.8018 | -9.3738 | -2.1215 | -7.3194 | -11.257 | -2.4941 | -7.1556
4.0642 | -9.5711 | -2.1898 | -7.4794 | -11.2444 | -2.9796 | -6.5609 | -9.5231 | -2.2708 | -7.164 | -11.4583 | -2.6953 | -6.8795
4.1341 | -9.6586 | -2.2773 | -7.2352 | -11.3881 | -3.1233 | -6.3536 | -9.6397 | -2.3874 | -6.9924 | -11.6111 | -2.8482 | -6.6269
4.204 | -9.7564 | -2.3751 | -7.053 | -11.5108 | -3.246 | -6.1388 | -9.7592 | -2.5069 | -6.8242 | -11.7717 | -3.0088 | -6.3578
4.2739 | -9.8476 | -2.4663 | -6.8245 | -11.6794 | -3.4146 | -5.9427 | -9.9066 | -2.6543 | -6.6574 | -11.9395 | -3.1765 | -6.1065
4.3438 | -9.9665 | -2.5852 | -6.6712 | -11.8102 | -3.5454 | -5.7495 | -10.0497 | -2.7974 | -6.5201 | -12.0982 | -3.3352 | -5.8633
4.4137 | -10.1239 | -2.7426 | -6.5067 | -11.9741 | -3.7093 | -5.5883 | -10.2156 | -2.9633 | -6.374 | -12.2964 | -3.5334 | -5.669
4.4836 | -10.2488 | -2.8675 | -6.343 | -12.1161 | -3.8513 | -5.4244 | -10.3181 | -3.0658 | -6.2209 | -12.4597 | -3.6968 | -5.4836
4.5535 | -10.3651 | -2.9838 | -6.2005 | -12.251 | -3.9863 | -5.293 | -10.4753 | -3.223 | -6.0799 | -12.6944 | -3.9314 | -5.3085
4.6234 | -10.4644 | -3.0831 | -6.0535 | -12.4016 | -4.1368 | -5.1657 | -10.5943 | -3.342 | -5.9287 | -12.8646 | -4.1016 | -5.1777
4.6933 | -10.5971 | -3.2158 | -5.9213 | -12.5423 | -4.2775 | -5.0571 | -10.7632 | -3.5109 | -5.762 | -13.0835 | -4.3206 | -5.0503
4.7632 | -10.7347 | -3.3534 | -5.7592 | -12.7075 | -4.4427 | -4.945 | -10.9078 | -3.6555 | -5.5998 | -13.2506 | -4.4876 | -4.9401
4.8331 | -10.8947 | -3.5134 | -5.5976 | -12.8718 | -4.607 | -4.8448 | -11.044 | -3.7917 | -5.4563 | -13.4394 | -4.6765 | -4.8261
4.903 | -11.048 | -3.6667 | -5.4341 | -13.0256 | -4.7608 | -4.7262 | -11.2032 | -3.9509 | -5.3285 | -13.6529 | -4.8899 | -4.7204
4.9729 | -11.2315 | -3.8503 | -5.3013 | -13.2227 | -4.9579 | -4.6088 | -11.3364 | -4.0841 | -5.1981 | -13.8652 | -5.1023 | -4.6189
5.0428 | -11.3982 | -4.0169 | -5.1648 | -13.4643 | -5.1995 | -4.5026 | -11.5237 | -4.2714 | -5.0678 | -14.092 | -5.3291 | -4.5238
5.1127 | -11.5667 | -4.1855 | -5.0307 | -13.7388 | -5.474 | -4.3935 | -11.6835 | -4.4312 | -4.9341 | -14.2593 | -5.4963 | -4.4308
5.1826 | -11.7357 | -4.3544 | -4.9084 | -13.9634 | -5.6986 | -4.2823 | -11.8963 | -4.644 | -4.8063 | -14.4759 | -5.713 | -4.3425
5.2525 | -11.9279 | -4.5466 | -4.8116 | -14.1867 | -5.9219 | -4.2026 | -12.0737 | -4.8214 | -4.6771 | -14.6818 | -5.9188 | -4.2678
5.3224 | -12.0727 | -4.6914 | -4.704 | -14.4076 | -6.1428 | -4.1259 | -12.2876 | -5.0353 | -4.5583 | -14.9283 | -6.1653 | -4.1899
5.3923 | -12.2353 | -4.854 | -4.596 | -14.6439 | -6.3791 | -4.0534 | -12.4699 | -5.2176 | -4.4376 | -15.1151 | -6.3522 | -4.1173
5.4622 | -12.4038 | -5.0225 | -4.5039 | -14.8629 | -6.5981 | -3.9803 | -12.7253 | -5.473 | -4.3192 | -15.3365 | -6.5736 | -4.0482
5.5321 | -12.5946 | -5.2133 | -4.4138 | -15.0854 | -6.8206 | -3.9154 | -12.9647 | -5.7125 | -4.2094 | -15.5747 | -6.8118 | -3.9775
5.602 | -12.8023 | -5.4211 | -4.2888 | -15.3382 | -7.0734 | -3.8536 | -13.2097 | -5.9574 | -4.086 | -15.7861 | -7.0232 | -3.9099
5.6719 | -13.0143 | -5.633 | -4.1683 | -15.5518 | -7.287 | -3.7926 | -13.4579 | -6.2056 | -3.972 | -16.0236 | -7.2607 | -3.8444
5.7418 | -13.2981 | -5.9168 | -4.0586 | -15.7742 | -7.5094 | -3.7388 | -13.6842 | -6.4319 | -3.8759 | -16.1927 | -7.4297 | -3.773
5.8117 | -13.5602 | -6.1789 | -3.9385 | -15.9941 | -7.7293 | -3.6795 | -13.9769 | -6.7246 | -3.7759 | -16.4396 | -7.6767 | -3.7103
5.8816 | -13.8623 | -6.481 | -3.8246 | -16.2738 | -8.009 | -3.6191 | -14.2345 | -6.9822 | -3.6791 | -16.6356 | -7.8726 | -3.6434
5.9515 | -14.1297 | -6.7484 | -3.7179 | -16.5102 | -8.2454 | -3.5446 | -14.5297 | -7.2774 | -3.5927 | -16.8735 | -8.1106 | -3.5651
6.0214 | -14.4588 | -7.0775 | -3.6194 | -16.79 | -8.5252 | -3.4744 | -14.7861 | -7.5338 | -3.5095 | -17.088 | -8.325 | -3.4946
6.0913 | -14.776 | -7.3947 | -3.5225 | -17.0182 | -8.7534 | -3.4135 | -15.0739 | -7.8216 | -3.4232 | -17.319 | -8.556 | -3.4314
6.1612 | -15.1206 | -7.7394 | -3.4545 | -17.2397 | -8.9749 | -3.3516 | -15.3943 | -8.142 | -3.3414 | -17.5431 | -8.7802 | -3.3622
6.2311 | -15.4088 | -8.0275 | -3.364 | -17.4458 | -9.181 | -3.2885 | -15.7467 | -8.4944 | -3.2561 | -17.7613 | -8.9984 | -3.305
6.301 | -15.764 | -8.3828 | -3.2785 | -17.6695 | -9.4047 | -3.227 | -16.1169 | -8.8646 | -3.1797 | -18.0142 | -9.2513 | -3.2475
6.3709 | -16.1751 | -8.7938 | -3.1947 | -18.0024 | -9.7376 | -3.1597 | -16.4993 | -9.247 | -3.0947 | -18.2751 | -9.5121 | -3.1863
6.4408 | -16.6543 | -9.2731 | -3.1106 | -18.3211 | -10.0563 | -3.0825 | -16.9505 | -9.6982 | -3.0069 | -18.5936 | -9.8307 | -3.1169
6.5107 | -17.0795 | -9.6982 | -3.0173 | -18.6736 | -10.4088 | -3.0106 | -17.4483 | -10.196 | -2.9269 | -18.9533 | -10.1904 | -3.0456
6.5806 | -17.4862 | -10.105 | -2.9394 | -18.9952 | -10.7304 | -2.9367 | -17.9723 | -10.72 | -2.8454 | -19.3672 | -10.6043 | -2.967
6.6505 | -17.9584 | -10.5771 | -2.8687 | -19.4018 | -11.137 | -2.8667 | -18.5106 | -11.2583 | -2.7638 | -19.8289 | -11.066 | -2.892
6.7204 | -18.4425 | -11.0612 | -2.8025 | -19.8173 | -11.5525 | -2.8053 | -19.0572 | -11.8049 | -2.691 | -20.3147 | -11.5518 | -2.8199
6.7903 | -18.9833 | -11.602 | -2.7385 | -20.2804 | -12.0156 | -2.7381 | -19.6727 | -12.4204 | -2.6294 | -20.9057 | -12.1427 | -2.7465
6.8602 | -19.4943 | -12.113 | -2.6722 | -20.72 | -12.4552 | -2.6748 | -20.2569 | -13.0046 | -2.5669 | -21.5373 | -12.7743 | -2.6699
6.9301 | -20.0472 | -12.666 | -2.6247 | -21.2208 | -12.956 | -2.6272 | -20.882 | -13.6297 | -2.5219 | -22.2552 | -13.4922 | -2.6113
7 | -20.4037 | -13.0224 | -2.5945 | -21.5683 | -13.3035 | -2.5937 | -21.2684 | -14.0161 | -2.4894 | -22.7251 | -13.9622 | -2.5677

## IP3
RF Frequency - GHz | Std "A" +20 | Output IP3 | Std "B" +20 | Output IP3
Average=> | 24.946 | 16.9733 | 23.4249 | 14.5612
1 | -1.79 | -76.9109 | 22.5232 | -3.8261
1.2602 | -1.1464 | -73.0328 | 22.9161 | -0.61816
1.5203 | -0.82767 | -64.6535 | 23.8928 | 2.5093
1.7805 | 0.58754 | -51.5414 | 24.0213 | 5.3491
2.0406 | 1.8713 | -37.0258 | 24.2444 | 6.457
2.3008 | 8.2545 | -18.0685 | 24.6794 | 7.5494
2.5609 | 13.4893 | -5.3244 | 25.3229 | 8.3949
2.8211 | 17.6841 | 3.2816 | 25.5537 | 10.0846
3.0812 | 17.6302 | 3.976 | 25.7411 | 11.6075
3.3414 | 16.9503 | 4.0182 | 25.645 | 13.6519
3.6015 | 19.2746 | 7.9179 | 25.3421 | 14.688
3.8617 | 19.9599 | 9.5852 | 23.8687 | 15.5764
4.1218 | 20.7444 | 11.2275 | 21.1363 | 14.846
4.382 | 19.7885 | 10.6234 | 19.2651 | 12.8337
4.6421 | 19.3311 | 10.6692 | 18.5321 | 11.5465
4.9023 | 20.0762 | 11.9867 | 19.5136 | 11.004
5.1624 | 20.7511 | 12.9773 | 21.3596 | 12.067
5.4226 | 23.1114 | 15.6977 | 22.5761 | 13.7076
5.6828 | 24.8108 | 17.6308 | 24.0509 | 14.7118
5.9429 | 25.6966 | 18.7019 | 25.1715 | 16.0247
6.2031 | 25.3598 | 18.4815 | 25.7992 | 17.0939
6.4632 | 25.9417 | 19.1912 | 24.0825 | 17.7571
6.7234 | 25.3555 | 18.6783 | 24.1676 | 16.066
6.9835 | 26.9579 | 20.3042 | 25.4442 | 16.1322
7.2437 | 26.4359 | 19.7598 | 27.9618 | 17.4514
7.5038 | 26.9268 | 20.1583 | 28.0961 | 19.966
7.764 | 25.5908 | 18.6823 | 26.6418 | 20.0362
8.0241 | 26.7691 | 19.671 | 26.4757 | 18.3755
8.2843 | 26.6882 | 19.4237 | 26.5991 | 17.9031
8.5444 | 27.1744 | 19.7344 | 26.9292 | 17.6534
8.8046 | 25.5071 | 17.9687 | 27.4367 | 17.6451
9.0647 | 27.2553 | 19.5411 | 26.0738 | 18.0438
9.3249 | 27.4026 | 19.5725 | 26.2067 | 16.8192
9.5851 | 27.6017 | 19.7092 | 25.9411 | 17.1991
9.8452 | 26.2491 | 18.3002 | 27.123 | 17.2184
10.1054 | 25.65 | 17.707 | 26.7242 | 18.4995
10.3655 | 25.1465 | 17.2128 | 25.8093 | 18.251
10.6257 | 24.4456 | 16.5055 | 24.5787 | 17.3617
10.8858 | 24.4362 | 16.5221 | 24.3994 | 16.2745
11.146 | 25.6969 | 17.7508 | 24.1586 | 16.1468
11.4061 | 26.0909 | 18.2437 | 24.1446 | 15.9748
11.6663 | 26.017 | 18.1323 | 23.7108 | 15.9561
11.9264 | 25.8692 | 17.8921 | 23.1589 | 15.4758
12.1866 | 25.8812 | 17.6985 | 22.724 | 14.7994
12.4467 | 26.7135 | 18.3359 | 22.8153 | 14.2263
12.7069 | 25.9995 | 17.5249 | 23.0709 | 14.1788
12.967 | 25.4027 | 17.0385 | 23.6104 | 14.3642
13.2272 | 25.3118 | 17.1836 | 23.7243 | 14.8816
13.4873 | 24.2543 | 16.2876 | 24.6055 | 14.9675
13.7475 | 23.6152 | 15.7738 | 25.6665 | 15.8368
14.0077 | 22.9691 | 15.1009 | 26.9864 | 16.8412
14.2678 | 24.1952 | 16.188 | 27.0664 | 18.0266
14.528 | 24.3487 | 16.2273 | 28.5288 | 17.9283
14.7881 | 25.2159 | 17.0877 | 28.3679 | 19.1337
15.0483 | 24.8697 | 16.8218 | 29.0602 | 18.8377
15.3084 | 26.2926 | 18.2377 | 27.3263 | 19.4791
15.5686 | 26.8837 | 18.8861 | 26.5609 | 17.7247
15.8287 | 26.1682 | 18.3158 | 26.6238 | 16.9739
16.0889 | 25.3895 | 17.5412 | 26.0228 | 16.8298
16.349 | 24.0563 | 16.3822 | 25.6713 | 16.3597
16.6092 | 23.904 | 16.3757 | 24.9144 | 16.0199
16.8693 | 23.1 | 15.7996 | 24.5219 | 15.611
17.1295 | 22.9472 | 15.6291 | 24.3111 | 15.3787
17.3896 | 23.0384 | 15.6437 | 23.4871 | 15.3296
17.6498 | 23.0357 | 15.6678 | 23.5098 | 14.6112
17.9099 | 23.2406 | 15.7699 | 23.515 | 14.5757
18.1701 | 23.2805 | 15.785 | 23.1554 | 14.5941
18.4303 | 23.7306 | 16.131 | 22.444 | 14.2255
18.6904 | 23.6354 | 16.0195 | 21.8058 | 13.5293
18.9506 | 24.231 | 16.5056 | 21.3624 | 12.9654
19.2107 | 24.3072 | 16.478 | 21.2443 | 12.4675
19.4709 | 25.177 | 17.236 | 20.9355 | 12.3115
19.731 | 25.2316 | 17.1705 | 20.761 | 11.8894
19.9912 | 25.3069 | 17.0506 | 20.3907 | 11.5761
20.2513 | 25.2806 | 16.7883 | 20.359 | 11.0559
20.5115 | 25.2411 | 16.5998 | 20.496 | 10.8827
20.7716 | 25.3681 | 16.627 | 20.5989 | 10.9236
21.0318 | 25.1349 | 16.1724 | 20.3527 | 10.9175
21.2919 | 25.1401 | 16.0798 | 20.1033 | 10.6337
21.5521 | 25.3166 | 16.148 | 19.9555 | 10.3735
21.8122 | 24.9251 | 15.6619 | 20.211 | 10.239
22.0724 | 24.5021 | 15.1431 | 20.1155 | 10.5053
22.3326 | 24.6834 | 15.324 | 20.0809 | 10.3917
22.5927 | 24.7685 | 15.5912 | 18.756 | 10.3488
22.8529 | 24.8823 | 15.8239 | 17.4755 | 8.8178
23.113 | 24.2645 | 15.1826 | 15.9482 | 7.3267
23.3732 | 24.6108 | 15.4252 | 15.6107 | 5.5694
23.6333 | 24.9711 | 15.5332 | 14.89 | 5.0687
23.8935 | 25.781 | 16.1936 | 14.2611 | 4.1975
24.1536 | 26.3666 | 16.644 | 13.9876 | 3.4079
24.4138 | 26.2847 | 16.4585 | 14.4594 | 3.0169
24.6739 | 25.4573 | 15.5466 | 15.7212 | 3.2559
24.9341 | 24.5725 | 14.5519 | 17.3018 | 4.237
25.1942 | 23.788 | 13.5695 | 20.091 | 5.4225
25.4544 | 23.4923 | 13.0771 | 21.8773 | 7.7647
25.7145 | 23.054 | 12.3405 | 24.3957 | 9.0166
25.9747 | 22.6623 | 11.7906 | 24.7909 | 10.8249
26.2348 | 22.4388 | 11.2504 | 25.5265 | 10.5323
26.495 | 22.4339 | 11.1081 | 22.5926 | 10.8018

## LO Harm-A
LO Output | 2xLO to IF | 2xLO to RF | LO Output | 3xLO to IF | 3xLO to RF | LO Output | 4xLO to IF | 4xLO to RF | LO Output | 5xLO to IF | 5xLO to RF | A Data
9 | -50.4587 | -47.2307 | 13.5 | -45.2931 | -58.7313 | 18 | -61.4455 | -47.859 | 22.5 | -56.2805 | -61.2486
9.3333 | -51.4409 | -46.3816 | 13.7396 | -45.3335 | -59.2612 | 18.1458 | -62.5888 | -47.6781 | 22.5521 | -56.3409 | -61.545
9.6667 | -52.6327 | -45.5072 | 13.9792 | -45.3988 | -60.0236 | 18.2917 | -63.7542 | -47.4175 | 22.6042 | -56.2656 | -62.0802
10 | -53.1601 | -44.1393 | 14.2188 | -46.4679 | -60.9992 | 18.4375 | -64.302 | -47.4489 | 22.6562 | -55.7349 | -62.5823
10.3333 | -54.157 | -42.9017 | 14.4583 | -47.0761 | -61.7065 | 18.5833 | -64.5615 | -47.0157 | 22.7083 | -55.3448 | -63.1999
10.6667 | -56.7775 | -41.4327 | 14.6979 | -47.7023 | -61.9676 | 18.7292 | -65.8292 | -46.8619 | 22.7604 | -55.1722 | -63.4497
11 | -61.1128 | -40.7055 | 14.9375 | -48.436 | -62.2954 | 18.875 | -66.0997 | -46.8388 | 22.8125 | -54.6022 | -63.0808
11.3333 | -68.5151 | -41.076 | 15.1771 | -48.9416 | -62.9983 | 19.0208 | -66.1822 | -46.9606 | 22.8646 | -54.8363 | -62.9436
11.6667 | -69.5487 | -41.6814 | 15.4167 | -49.4649 | -63.3091 | 19.1667 | -64.8957 | -46.5956 | 22.9167 | -54.8959 | -63.4053
12 | -67.3846 | -42.6449 | 15.6562 | -48.0547 | -63.9551 | 19.3125 | -65.2813 | -46.4266 | 22.9688 | -55.1594 | -64.3828
12.3333 | -60.1162 | -43.1352 | 15.8958 | -47.2379 | -63.5124 | 19.4583 | -64.7131 | -46.4669 | 23.0208 | -55.0103 | -65.5387
12.6667 | -57.2975 | -43.8415 | 16.1354 | -45.6789 | -63.491 | 19.6042 | -64.13 | -46.8466 | 23.0729 | -54.5046 | -65.7958
13 | -55.4196 | -44.4189 | 16.375 | -44.1165 | -62.7119 | 19.75 | -63.2632 | -46.9167 | 23.125 | -54.1636 | -66.1467
13.3333 | -54.4732 | -44.9193 | 16.6146 | -42.561 | -62.4647 | 19.8958 | -63.5049 | -46.7717 | 23.1771 | -53.8173 | -66.0537
13.6667 | -54.7203 | -45.4893 | 16.8542 | -42.4524 | -62.0622 | 20.0417 | -63.583 | -46.5358 | 23.2292 | -53.8609 | -66.646
14 | -55.2551 | -46.1605 | 17.0938 | -43.2293 | -61.4023 | 20.1875 | -63.4817 | -46.3576 | 23.2812 | -53.8964 | -66.8849
14.3333 | -56.1975 | -46.9453 | 17.3333 | -44.9404 | -60.6447 | 20.3333 | -63.2143 | -46.0988 | 23.3333 | -53.9983 | -67.2893
14.6667 | -55.4371 | -47.6872 | 17.5729 | -46.1249 | -60.0208 | 20.4792 | -64.1226 | -45.964 | 23.3854 | -54.118 | -68.1652
15 | -55.026 | -48.4109 | 17.8125 | -48.0169 | -59.083 | 20.625 | -65.3387 | -45.8234 | 23.4375 | -54.0519 | -69.5159
15.3333 | -54.6698 | -49.2892 | 18.0521 | -49.7131 | -57.6998 | 20.7708 | -66.2109 | -45.6197 | 23.4896 | -53.9884 | -70.4735
15.6667 | -54.8548 | -50.8253 | 18.2917 | -50.9222 | -56.775 | 20.9167 | -66.8375 | -45.3498 | 23.5417 | -53.4888 | -71.3285
16 | -55.6731 | -53.2237 | 18.5312 | -51.3822 | -56.2189 | 21.0625 | -68.7294 | -45.0582 | 23.5938 | -53.2291 | -71.7915
16.3333 | -56.7288 | -58.3941 | 18.7708 | -51.1099 | -56.1585 | 21.2083 | -70.9897 | -44.8613 | 23.6458 | -52.7728 | -71.5816
16.6667 | -59.7796 | -65.198 | 19.0104 | -51.2723 | -56.3518 | 21.3542 | -72.365 | -44.6221 | 23.6979 | -52.8201 | -70.9536
17 | -61.3378 | -66.9085 | 19.25 | -52.208 | -56.9157 | 21.5 | -72.901 | -44.4351 | 23.75 | -52.3504 | -69.9646
17.3333 | -64.4361 | -63.8456 | 19.4896 | -53.1568 | -58.0821 | 21.6458 | -74.6887 | -44.166 | 23.8021 | -52.2969 | -70.2826
17.6667 | -67.415 | -57.1666 | 19.7292 | -54.296 | -58.5316 | 21.7917 | -74.868 | -43.8747 | 23.8542 | -51.8007 | -70.5921
18 | -68.3282 | -53.526 | 19.9688 | -54.4279 | -58.8984 | 21.9375 | -80.1731 | -43.6115 | 23.9062 | -51.9655 | -71.506
18.3333 | -65.8311 | -51.1916 | 20.2083 | -55.1514 | -58.5084 | 22.0833 | -81.9389 | -43.6654 | 23.9583 | -52.0783 | -71.6886
18.6667 | -61.8867 | -49.5719 | 20.4479 | -56.3546 | -58.5393 | 22.2292 | -82.3919 | -44.1901 | 24.0104 | -52.0655 | -71.1299
19 | -59.4416 | -48.1873 | 20.6875 | -57.7029 | -58.4392 | 22.375 | -78.7897 | -44.6433 | 24.0625 | -51.2462 | -70.1657
19.3333 | -58.3211 | -46.7149 | 20.9271 | -59.1224 | -58.9473 | 22.5208 | -75.0635 | -44.6679 | 24.1146 | -50.0837 | -69.2911
19.6667 | -56.8965 | -45.2534 | 21.1667 | -63.3533 | -59.5112 | 22.6667 | -75.4804 | -44.6569 | 24.1667 | -49.6835 | -69.2482
20 | -56.5768 | -44.1844 | 21.4062 | -66.6274 | -59.9596 | 22.8125 | -74.1047 | -45.2595 | 24.2188 | -49.8403 | -69.1285
20.3333 | -56.1662 | -43.7598 | 21.6458 | -67.0664 | -59.8608 | 22.9583 | -75.0469 | -46.0111 | 24.2708 | -49.9419 | -69.0547
20.6667 | -55.2343 | -43.7933 | 21.8854 | -62.1185 | -60.0909 | 23.1042 | -75.0902 | -46.4234 | 24.3229 | -49.6233 | -69.3115
21 | -54.8224 | -43.972 | 22.125 | -57.0898 | -59.4712 | 23.25 | -74.282 | -46.404 | 24.375 | -49.3757 | -69.5572
21.3333 | -54.5939 | -43.7739 | 22.3646 | -53.2285 | -58.7099 | 23.3958 | -73.4708 | -46.8305 | 24.4271 | -49.6259 | -69.8066
21.6667 | -54.8982 | -43.2845 | 22.6042 | -50.5133 | -57.3915 | 23.5417 | -71.0962 | -46.9839 | 24.4792 | -50.3464 | -69.0949
22 | -54.5524 | -43.0292 | 22.8438 | -48.5933 | -56.9104 | 23.6875 | -70.7685 | -47.5115 | 24.5312 | -50.5603 | -68.6059
22.3333 | -54.2184 | -43.1749 | 23.0833 | -47.0825 | -57.1869 | 23.8333 | -69.1468 | -47.8042 | 24.5833 | -50.0938 | -68.0364
22.6667 | -53.6846 | -43.6868 | 23.3229 | -45.8684 | -57.2627 | 23.9792 | -70.0361 | -49.027 | 24.6354 | -49.13 | -68.1966
23 | -53.6555 | -43.9417 | 23.5625 | -44.2831 | -57.9957 | 24.125 | -69.2751 | -49.5202 | 24.6875 | -48.4391 | -67.7687
23.3333 | -53.7377 | -43.892 | 23.8021 | -43.234 | -58.4972 | 24.2708 | -69.4376 | -49.8059 | 24.7396 | -48.2544 | -68.0052
23.6667 | -54.044 | -43.5512 | 24.0417 | -42.4536 | -59.1586 | 24.4167 | -69.0358 | -49.9875 | 24.7917 | -48.1961 | -67.5534
24 | -53.8688 | -43.1849 | 24.2812 | -42.366 | -59.7455 | 24.5625 | -69.2111 | -51.3849 | 24.8438 | -48.0363 | -67.5652
24.3333 | -53.9542 | -42.7356 | 24.5208 | -41.9955 | -61.1067 | 24.7083 | -69.0903 | -52.7231 | 24.8958 | -48.0559 | -66.3693
24.6667 | -54.2987 | -42.5271 | 24.7604 | -41.7421 | -61.7308 | 24.8542 | -69.0831 | -53.3315 | 24.9479 | -47.8536 | -66.0925
25 | -54.8496 | -42.398 | 25 | -41.4229 | -61.9258 | 25 | -69.0216 | -53.0494 | 25 | -47.7673 | -66.0296

## LO Harm-B
LO Output | 2xLO to IF | 2xLO to RF | LO Output | 3xLO to IF | 3xLO to RF | LO Output | 4xLO to IF | 4xLO to RF | LO Output | 5xLO to IF | 5xLO to RF | B Data
9 | -46.372 | -27.6769 | 13.5 | -48.6679 | -54.3606 | 18 | -60.3736 | -40.5406 | 22.5 | -47.4568 | -63.8578
9.3333 | -46.597 | -27.2594 | 13.7396 | -47.3334 | -53.8348 | 18.1458 | -61.8244 | -41.0851 | 22.5521 | -46.7223 | -64.5778
9.6667 | -47.3743 | -26.904 | 13.9792 | -45.3367 | -53.2208 | 18.2917 | -62.8543 | -41.4187 | 22.6042 | -45.984 | -65.5348
10 | -48.4763 | -26.5196 | 14.2188 | -43.6234 | -52.572 | 18.4375 | -63.3171 | -41.3706 | 22.6562 | -45.7857 | -66.4506
10.3333 | -50.1286 | -26.3148 | 14.4583 | -42.0045 | -52.0727 | 18.5833 | -64.4874 | -41.9204 | 22.7083 | -45.7998 | -65.8844
10.6667 | -52.1259 | -25.9029 | 14.6979 | -40.5118 | -51.4772 | 18.7292 | -65.9682 | -42.1773 | 22.7604 | -45.4672 | -66.317
11 | -54.9708 | -25.6374 | 14.9375 | -39.2108 | -51.1292 | 18.875 | -66.6744 | -42.4266 | 22.8125 | -44.9825 | -67.4696
11.3333 | -58.93 | -25.3123 | 15.1771 | -37.9639 | -50.872 | 19.0208 | -66.3761 | -42.0884 | 22.8646 | -44.4794 | -73.0644
11.6667 | -64.3102 | -24.9212 | 15.4167 | -36.8425 | -50.8581 | 19.1667 | -66.1875 | -42.5994 | 22.9167 | -44.8028 | -71.6106
12 | -66.0028 | -24.6005 | 15.6562 | -36.1728 | -50.7219 | 19.3125 | -64.6165 | -42.8447 | 22.9688 | -45.4288 | -70.1024
12.3333 | -64.3298 | -24.2601 | 15.8958 | -35.4284 | -50.3497 | 19.4583 | -64.0926 | -43.0079 | 23.0208 | -45.7483 | -66.8677
12.6667 | -59.9062 | -24.2004 | 16.1354 | -34.8704 | -49.5916 | 19.6042 | -63.5216 | -43.1054 | 23.0729 | -45.4057 | -68.0785
13 | -57.522 | -24.2177 | 16.375 | -34.322 | -48.8277 | 19.75 | -63.7015 | -44.0684 | 23.125 | -44.5978 | -67.8216
13.3333 | -56.4953 | -24.4331 | 16.6146 | -33.9594 | -48.0608 | 19.8958 | -63.271 | -45.3459 | 23.1771 | -44.0941 | -64.8355
13.6667 | -56.2146 | -24.6605 | 16.8542 | -33.8565 | -47.7691 | 20.0417 | -62.396 | -47.3533 | 23.2292 | -43.9097 | -63.5353
14 | -57.0529 | -24.9208 | 17.0938 | -33.6698 | -47.6406 | 20.1875 | -61.7633 | -50.2728 | 23.2812 | -43.8529 | -62.5887
14.3333 | -58.3982 | -25.1241 | 17.3333 | -33.6018 | -47.5826 | 20.3333 | -60.9992 | -52.1684 | 23.3333 | -43.7176 | -62.6628
14.6667 | -60.4705 | -25.222 | 17.5729 | -33.5745 | -47.3623 | 20.4792 | -60.2333 | -53.4887 | 23.3854 | -43.2003 | -62.389
15 | -63.571 | -25.1768 | 17.8125 | -33.6225 | -46.923 | 20.625 | -59.8826 | -51.5285 | 23.4375 | -43.4951 | -64.2269
15.3333 | -65.1703 | -25.0712 | 18.0521 | -33.644 | -46.4461 | 20.7708 | -59.1219 | -49.2331 | 23.4896 | -43.5701 | -64.9698
15.6667 | -63.187 | -24.9881 | 18.2917 | -33.4853 | -46.2215 | 20.9167 | -57.9694 | -45.547 | 23.5417 | -43.678 | -64.9796
16 | -57.7163 | -24.8986 | 18.5312 | -33.2945 | -46.1197 | 21.0625 | -57.1988 | -43.1467 | 23.5938 | -42.9615 | -63.1262
16.3333 | -52.713 | -24.7991 | 18.7708 | -33.2936 | -46.1857 | 21.2083 | -55.7446 | -41.5046 | 23.6458 | -42.3183 | -60.9952
16.6667 | -49.2607 | -24.7715 | 19.0104 | -33.6118 | -46.2427 | 21.3542 | -55.1418 | -40.3579 | 23.6979 | -41.9983 | -60.1476
17 | -47.3127 | -24.8254 | 19.25 | -33.9168 | -46.3863 | 21.5 | -53.607 | -39.0043 | 23.75 | -41.8381 | -59.2324
17.3333 | -45.7943 | -24.9462 | 19.4896 | -34.4075 | -47.2942 | 21.6458 | -53.1587 | -37.9159 | 23.8021 | -41.7746 | -59.4224
17.6667 | -44.9736 | -24.9649 | 19.7292 | -34.8508 | -48.5983 | 21.7917 | -52.7158 | -37.1916 | 23.8542 | -41.7263 | -60.2115
18 | -44.1726 | -24.8443 | 19.9688 | -35.3472 | -51.0448 | 21.9375 | -52.3181 | -36.7139 | 23.9062 | -42.0376 | -62.0211
18.3333 | -43.5837 | -24.6561 | 20.2083 | -35.6111 | -52.8013 | 22.0833 | -51.9015 | -36.0336 | 23.9583 | -42.3447 | -63.2627
18.6667 | -43.0758 | -24.3526 | 20.4479 | -35.9123 | -56.4522 | 22.2292 | -51.3085 | -35.2059 | 24.0104 | -42.7985 | -62.9494
19 | -42.9934 | -24.177 | 20.6875 | -36.2258 | -61.8502 | 22.375 | -50.8316 | -34.6487 | 24.0625 | -42.2462 | -60.6019
19.3333 | -43.353 | -24.0314 | 20.9271 | -36.6801 | -67.639 | 22.5208 | -50.17 | -34.4043 | 24.1146 | -41.9023 | -58.6945
19.6667 | -44.3033 | -24.0066 | 21.1667 | -36.664 | -66.9179 | 22.6667 | -49.5596 | -34.3665 | 24.1667 | -41.2851 | -57.5577
20 | -45.4704 | -24.1042 | 21.4062 | -36.539 | -62.6582 | 22.8125 | -49.244 | -34.0012 | 24.2188 | -41.3453 | -58.0066
20.3333 | -46.4256 | -24.1892 | 21.6458 | -36.2355 | -57.9681 | 22.9583 | -49.2213 | -33.656 | 24.2708 | -41.117 | -58.4542
20.6667 | -47.0907 | -24.5004 | 21.8854 | -36.0073 | -56.4162 | 23.1042 | -48.748 | -33.0596 | 24.3229 | -41.228 | -58.4729
21 | -48.7289 | -24.8681 | 22.125 | -35.4907 | -55.3389 | 23.25 | -48.1446 | -32.7454 | 24.375 | -41.039 | -58.1742
21.3333 | -50.103 | -25.4534 | 22.3646 | -34.7331 | -54.5559 | 23.3958 | -47.6452 | -32.3962 | 24.4271 | -41.1431 | -58.4806
21.6667 | -51.3342 | -25.9094 | 22.6042 | -33.9686 | -53.9726 | 23.5417 | -47.6947 | -32.4586 | 24.4792 | -40.8774 | -58.7662
22 | -51.1164 | -26.3366 | 22.8438 | -33.336 | -53.5107 | 23.6875 | -47.5893 | -32.4817 | 24.5312 | -40.8534 | -58.7915
22.3333 | -50.7576 | -26.7769 | 23.0833 | -32.6567 | -52.0817 | 23.8333 | -46.8562 | -32.3019 | 24.5833 | -41.3101 | -58.4427
22.6667 | -50.5242 | -27.3896 | 23.3229 | -31.9791 | -51.1136 | 23.9792 | -46.217 | -32.0793 | 24.6354 | -41.7182 | -58.0171
23 | -49.937 | -27.9641 | 23.5625 | -30.9107 | -49.9142 | 24.125 | -45.8104 | -31.9398 | 24.6875 | -41.7267 | -57.8734
23.3333 | -49.4983 | -28.5275 | 23.8021 | -30.1378 | -49.0291 | 24.2708 | -45.4889 | -32.1102 | 24.7396 | -41.5596 | -57.3485
23.6667 | -48.5339 | -29.0068 | 24.0417 | -29.3797 | -48.2857 | 24.4167 | -45.4031 | -32.2427 | 24.7917 | -41.3507 | -56.7817
24 | -47.6573 | -29.6676 | 24.2812 | -29.0907 | -47.6095 | 24.5625 | -45.2966 | -32.36 | 24.8438 | -41.6181 | -56.8117
24.3333 | -47.0807 | -30.571 | 24.5208 | -28.6673 | -47.221 | 24.7083 | -45.5179 | -32.4998 | 24.8958 | -41.6123 | -57.3299
24.6667 | -46.5702 | -31.614 | 24.7604 | -28.5269 | -46.7257 | 24.8542 | -45.3906 | -32.6991 | 24.9479 | -41.305 | -57.8239
25 | -46.3737 | -32.315 | 25 | -28.4417 | -46.422 | 25 | -45.2975 | -32.9058 | 25 | -40.9626 | -57.9045

## 2Rx2L
RF Frequency | "A" 2RFx2LO -10 dBm | "A" 2RFx2LO -5 dBm | "B" 2RFx2LO -10 dBm | "B" 2RFx2LO -5 dBm
Average=> | -68.7851 | Actual Data | -67.2967 | Actual Data
4.5 | -66.0577 | -61.0577 | -55.1543 | -61.2326
4.7092 | -64.9341 | -59.9341 | -53.277 | -58.8397
4.9184 | -63.9941 | -58.9941 | -52.0532 | -55.6053
5.1276 | -63.4099 | -58.4099 | -53.8024 | -52.8549
5.3367 | -63.6885 | -58.6885 | -55.8597 | -51.0526
5.5459 | -65.5373 | -60.5373 | -57.7754 | -50.7384
5.7551 | -67.9962 | -62.9962 | -58.0384 | -51.4658
5.9643 | -71.6437 | -66.6437 | -59.6746 | -53.2646
6.1735 | -73.3962 | -68.3962 | -61.9002 | -55.4826
6.3827 | -78.379 | -73.379 | -65.3995 | -56.9734
6.5918 | -78.3408 | -73.3408 | -68.5847 | -59.1783
6.801 | -77.1505 | -72.1505 | -70.9024 | -63.0099
7.0102 | -71.9145 | -66.9145 | -70.1863 | -68.1064
7.2194 | -70.4098 | -65.4098 | -68.1122 | -72.2964
7.4286 | -68.8415 | -63.8415 | -65.1678 | -79.2024
7.6378 | -66.7065 | -61.7065 | -62.5198 | -78.0025
7.8469 | -65.2304 | -60.2304 | -60.7169 | -78.8119
8.0561 | -64.3632 | -59.3632 | -59.8914 | -73.1084
8.2653 | -63.4753 | -58.4753 | -59.5056 | -77.2641
8.4745 | -63.6439 | -58.6439 | -59.7734 | -73.9106
8.6837 | -64.6143 | -59.6143 | -59.5177 | -70.9948
8.8929 | -65.9788 | -60.9788 | -59.5967 | -65.9324
9.102 | -65.5874 | -60.5874 | -59.3022 | -64.424
9.3112 | -65.409 | -60.409 | -59.6732 | -62.864
9.5204 | -65.0459 | -60.0459 | -60.9601 | -60.8407
9.7296 | -65.6117 | -60.6117 | -62.4461 | -59.5494
9.9388 | -66.5565 | -61.5565 | -64.2377 | -58.8758
10.148 | -68.7595 | -63.7595 | -65.1806 | -58.5206
10.3571 | -69.7348 | -64.7348 | -64.9949 | -58.1175
10.5663 | -69.2602 | -64.2602 | -63.782 | -58.0914
10.7755 | -67.9817 | -62.9817 | -62.4517 | -58.3998
10.9847 | -66.4227 | -61.4227 | -61.842 | -59.0788
11.1939 | -65.2109 | -60.2109 | -60.9603 | -59.9634
11.4031 | -63.3564 | -58.3564 | -59.6729 | -61.3884
11.6122 | -62.7458 | -57.7458 | -58.7206 | -62.8268
11.8214 | -62.7198 | -57.7198 | -58.3428 | -63.9275
12.0306 | -62.4742 | -57.4742 | -58.8998 | -63.9236
12.2398 | -62.6405 | -57.6405 | -59.1778 | -64.0564
12.449 | -63.7364 | -58.7364 | -60.6023 | -64.7607
12.6582 | -67.5813 | -62.5813 | -63.8304 | -65.4469
12.8673 | -73.9037 | -68.9037 | -68.4864 | -65.2491
13.0765 | -77.7306 | -72.7306 | -71.1381 | -64.3017
13.2857 | -77.851 | -72.851 | -72.1613 | -63.717
13.4949 | -73.789 | -68.789 | -71.7175 | -64.2674
13.7041 | -70.7484 | -65.7484 | -76.4477 | -64.8703
13.9133 | -68.177 | -63.177 | -74.7843 | -65.3526
14.1224 | -66.7361 | -61.7361 | -72.6497 | -64.0697
14.3316 | -65.3895 | -60.3895 | -65.183 | -65.1911
14.5408 | -69.0114 | -64.0114 | -64.8309 | -67.8864
14.75 | -70.739 | -65.739 | -66.4979 | -71.8876
14.9592 | -73.1397 | -68.1397 | -72.71 | -72.7191
15.1684 | -70.9253 | -65.9253 | -77.7128 | -72.8287
15.3776 | -70.1636 | -65.1636 | -80.333 | -72.8245
15.5867 | -69.1058 | -64.1058 | -79.6398 | -73.3718
15.7959 | -68.5671 | -63.5671 | -76.3692 | -72.0676
16.0051 | -69.3738 | -64.3738 | -73.8687 | -70.6796
16.2143 | -70.5421 | -65.5421 | -70.5536 | -69.0513
16.4235 | -71.266 | -66.266 | -71.1803 | -68.3492
16.6327 | -71.8646 | -66.8646 | -73.3088 | -69.1997
16.8418 | -72.2552 | -67.2552 | -74.1454 | -75.8149
17.051 | -74.2677 | -69.2677 | -73.3518 | -82.2156
17.2602 | -76.0659 | -71.0659 | -71.3445 | -83.0134
17.4694 | -76.0492 | -71.0492 | -70.338 | -76.9112
17.6786 | -74.6094 | -69.6094 | -69.3834 | -71.9889
17.8878 | -72.666 | -67.666 | -69.4149 | -70.8709
18.0969 | -72.501 | -67.501 | -70.6796 | -71.4554
18.3061 | -72.5944 | -67.5944 | -71.4247 | -73.6486
18.5153 | -72.1507 | -67.1507 | -71.4653 | -73.988
18.7245 | -71.5293 | -66.5293 | -71.6597 | -71.822
18.9337 | -71.0482 | -66.0482 | -72.1984 | -67.3801
19.1429 | -71.4053 | -66.4053 | -72.2785 | -63.308
19.352 | -71.7626 | -66.7626 | -73.4829 | -61.0461
19.5612 | -72.6205 | -67.6205 | -75.2438 | -59.786
19.7704 | -73.0597 | -68.0597 | -76.961 | -62.5194
19.9796 | -74.0929 | -69.0929 | -84.6668 | -66.4125
20.1888 | -72.3903 | -67.3903 | -85.7335 | -68.8307
20.398 | -70.0923 | -65.0923 | -86.7048 | -69.139
20.6071 | -67.5414 | -62.5414 | -77.9263 | -67.9891
20.8163 | -67.1072 | -62.1073 | -75.0552 | -67.6836
21.0255 | -68.7297 | -63.7297 | -73.9627 | -66.3765
21.2347 | -69.4841 | -64.4841 | -70.6096 | -65.5401
21.4439 | -70.936 | -65.936 | -67.474 | -64.6468
21.6531 | -70.2602 | -65.2602 | -63.7721 | -64.8739
21.8622 | -70.3355 | -65.3355 | -62.8353 | -64.6077
22.0714 | -69.5189 | -64.5189 | -62.9903 | -64.4818
22.2806 | -68.8462 | -63.8462 | -63.7046 | -64.3869
22.4898 | -66.8582 | -61.8582 | -66.0659 | -65.7236
22.699 | -65.0539 | -60.0539 | -68.3867 | -65.9031
22.9082 | -63.1326 | -58.1326 | -69.8719 | -66.04
23.1173 | -62.2124 | -57.2124 | -70.0197 | -64.3959
23.3265 | -61.2701 | -56.2701 | -69.6771 | -58.6321
23.5357 | -61.1633 | -56.1633 | -69.4539 | -52.8042
23.7449 | -61.1002 | -56.1002 | -67.4452 | -48.3346
23.9541 | -61.5222 | -56.5222 | -64.0996 | -48.6754
24.1633 | -61.998 | -56.998 | -60.6943 | -50.1028
24.3724 | -62.5386 | -57.5386 | -58.9593 | -52.1981
24.5816 | -62.5182 | -57.5182 | -57.711 | -53.6748
24.7908 | -62.6751 | -57.6751 | -56.2675 | -53.5521
25 | -62.8198 | -57.8198 | -54.4446 | -52.6376

## 2Ix1L
RF Frequency | "A" 2IFx1LO -10 dBm | "A" 2IFx1LO -5 dBm | "B" 2IFx1LO -10 dBm | "B" 2IFx1LO -5 dBm
Average=> | -70.2437 | Actual Data | -70.4161 | Actual Data
4.5 | -67.064 | -62.064 | -66.1247 | -61.1247
4.7092 | -66.7346 | -61.7346 | -64.969 | -59.969
4.9184 | -66.4956 | -61.4956 | -65.9444 | -60.9444
5.1276 | -64.9612 | -59.9612 | -66.7713 | -61.7713
5.3367 | -64.8698 | -59.8698 | -69.1087 | -64.1087
5.5459 | -66.8119 | -61.8119 | -71.0481 | -66.0481
5.7551 | -69.5941 | -64.5941 | -71.0881 | -66.0881
5.9643 | -73.8198 | -68.8198 | -72.1415 | -67.1415
6.1735 | -74.874 | -69.874 | -72.9859 | -67.9859
6.3827 | -77.3684 | -72.3684 | -75.0529 | -70.0529
6.5918 | -78.5272 | -73.5272 | -79.0128 | -74.0128
6.801 | -79.5805 | -74.5805 | -81.4752 | -76.4752
7.0102 | -80.2221 | -75.2221 | -83.9959 | -78.9959
7.2194 | -80.7101 | -75.7101 | -84.9595 | -79.9595
7.4286 | -81.1738 | -76.1738 | -82.6617 | -77.6617
7.6378 | -80.1691 | -75.1691 | -81.4096 | -76.4096
7.8469 | -77.0513 | -72.0513 | -78.5608 | -73.5608
8.0561 | -75.9219 | -70.9219 | -78.5331 | -73.5331
8.2653 | -77.3084 | -72.3084 | -76.5413 | -71.5413
8.4745 | -76.2396 | -71.2396 | -73.6083 | -68.6083
8.6837 | -72.2968 | -67.2968 | -70.6337 | -65.6337
8.8929 | -67.1868 | -62.1868 | -68.128 | -63.128
9.102 | -71.0026 | -66.0026 | -67.4319 | -62.4319
9.3112 | -74.9297 | -69.9297 | -68.5012 | -63.5012
9.5204 | -77.1019 | -72.1019 | -72.424 | -67.424
9.7296 | -73.124 | -68.124 | -74.5452 | -69.5452
9.9388 | -69.1672 | -64.1672 | -75.0542 | -70.0542
10.148 | -66.8395 | -61.8395 | -72.4327 | -67.4327
10.3571 | -68.938 | -63.938 | -72.6142 | -67.6142
10.5663 | -74.6401 | -69.6401 | -74.2397 | -69.2397
10.7755 | -80.6689 | -75.6689 | -77.1003 | -72.1003
10.9847 | -83.0046 | -78.0046 | -79.1369 | -74.1369
11.1939 | -79.3708 | -74.3708 | -79.7439 | -74.7439
11.4031 | -74.6247 | -69.6247 | -78.6716 | -73.6716
11.6122 | -70.8041 | -65.8041 | -76.1739 | -71.1739
11.8214 | -75.4021 | -70.4021 | -73.5036 | -68.5036
12.0306 | -80.781 | -75.781 | -71.224 | -66.224
12.2398 | -85.7304 | -80.7304 | -71.6112 | -66.6112
12.449 | -84.0182 | -79.0182 | -71.3588 | -66.3588
12.6582 | -80.0982 | -75.0982 | -71.8985 | -66.8985
12.8673 | -74.6757 | -69.6757 | -72.2838 | -67.2838
13.0765 | -71.314 | -66.314 | -77.0889 | -72.0889
13.2857 | -69.0713 | -64.0713 | -78.5615 | -73.5615
13.4949 | -67.745 | -62.745 | -79.636 | -74.636
13.7041 | -66.9721 | -61.9721 | -75.3549 | -70.3549
13.9133 | -66.1535 | -61.1535 | -73.4021 | -68.4021
14.1224 | -66.2105 | -61.2105 | -71.0873 | -66.0873
14.3316 | -63.8782 | -58.8782 | -73.9229 | -68.9229
14.5408 | -63.1083 | -58.1083 | -78.842 | -73.842
14.75 | -64.0731 | -59.0731 | -78.678 | -73.678
14.9592 | -67.95 | -62.95 | -73.7383 | -68.7383
15.1684 | -69.1901 | -64.1901 | -66.0549 | -61.0549
15.3776 | -66.708 | -61.708 | -62.0595 | -57.0595
15.5867 | -63.4056 | -58.4056 | -59.1833 | -54.1833
15.7959 | -62.9432 | -57.9432 | -58.2088 | -53.2088
16.0051 | -63.6332 | -58.6332 | -58.3588 | -53.3588
16.2143 | -67.56 | -62.56 | -60.3488 | -55.3488
16.4235 | -73.2242 | -68.2242 | -63.3527 | -58.3527
16.6327 | -79.0198 | -74.0198 | -66.8462 | -61.8462
16.8418 | -81.4218 | -76.4218 | -68.9108 | -63.9108
17.051 | -79.3889 | -74.3889 | -69.1519 | -64.1519
17.2602 | -77.0674 | -72.0674 | -68.8954 | -63.8954
17.4694 | -75.1146 | -70.1146 | -69.3742 | -64.3742
17.6786 | -73.6588 | -68.6588 | -71.4354 | -66.4354
17.8878 | -72.141 | -67.141 | -73.6774 | -68.6774
18.0969 | -70.2011 | -65.2011 | -77.6929 | -72.6929
18.3061 | -70.0138 | -65.0138 | -77.536 | -72.536
18.5153 | -71.0649 | -66.0649 | -75.9707 | -70.9707
18.7245 | -74.0802 | -69.0802 | -71.2201 | -66.2201
18.9337 | -74.4755 | -69.4755 | -68.2048 | -63.2048
19.1429 | -72.8848 | -67.8848 | -66.0962 | -61.0962
19.352 | -68.9942 | -63.9942 | -64.7899 | -59.7899
19.5612 | -66.3017 | -61.3017 | -64.2946 | -59.2946
19.7704 | -63.564 | -58.564 | -63.0832 | -58.0832
19.9796 | -60.4853 | -55.4853 | -60.8337 | -55.8337
20.1888 | -58.2357 | -53.2357 | -58.5356 | -53.5356
20.398 | -58.556 | -53.556 | -57.5048 | -52.5048
20.6071 | -61.6735 | -56.6735 | -57.9297 | -52.9297
20.8163 | -65.5775 | -60.5775 | -58.7762 | -53.7762
21.0255 | -69.4014 | -64.4014 | -59.7085 | -54.7085
21.2347 | -71.2674 | -66.2674 | -61.1955 | -56.1955
21.4439 | -70.8478 | -65.8478 | -63.1415 | -58.1415
21.6531 | -67.8642 | -62.8642 | -66.163 | -61.163
21.8622 | -64.0965 | -59.0965 | -69.4456 | -64.4456
22.0714 | -61.5905 | -56.5905 | -70.5547 | -65.5547
22.2806 | -60.0727 | -55.0727 | -71.1517 | -66.1517
22.4898 | -58.842 | -53.842 | -68.9544 | -63.9544
22.699 | -57.6885 | -52.6885 | -68.2872 | -63.2872
22.9082 | -57.3005 | -52.3005 | -66.4061 | -61.4061
23.1173 | -57.5317 | -52.5317 | -65.4278 | -60.4278
23.3265 | -57.8797 | -52.8797 | -64.5262 | -59.5262
23.5357 | -57.9218 | -52.9218 | -63.8087 | -58.8087
23.7449 | -58.8099 | -53.8099 | -63.6366 | -58.6366
23.9541 | -60.8295 | -55.8295 | -63.3592 | -58.3592
24.1633 | -64.3256 | -59.3256 | -63.1027 | -58.1027
24.3724 | -66.7414 | -61.7414 | -62.3471 | -57.3471
24.5816 | -69.049 | -64.049 | -61.8579 | -56.8579
24.7908 | -70.4835 | -65.4835 | -61.6978 | -56.6978
25 | -72.3005 | -67.3005 | -62.1483 | -57.1483